{"text_format":"5","workbook":"ods","n_sheets":18,"sheets":["Planilha1","PIRXIII","PIR001"," PIR003","PIR005","PIR007","PIR015","PIR017","PIR019","PIR024","PIR025","PIR027","PIR028","PIR029","PIR030","PIR031","PIR033","COMP. OD BANCADAxCTD"],"defined_names":["Excel_BuiltIn__FilterDatabase_1"],"not_rendered":{"charts":15,"images":16,"defined_names":1},"read":"whole","wrapped":"gz","date_system":"iso_in_file"}

{"sheet":"Planilha1","index":0,"n_rows":509,"n_cols":15,"rows_kept":509,"merged":[[0,0,0,14],[1,0,1,14],[3,0,3,14],[7,0,7,2],[8,0,8,2],[12,0,12,1],[24,0,24,1],[36,0,36,1],[48,0,48,1],[60,0,60,1],[72,0,72,1],[84,0,84,1],[96,0,96,1],[108,0,108,1],[120,0,120,1],[132,0,132,1],[144,0,144,1],[156,0,156,1],[168,0,168,1],[180,0,180,1],[192,0,192,1],[204,0,204,1],[216,0,216,1],[228,0,228,1],[240,0,240,1],[252,0,252,1],[264,0,264,1],[276,0,276,1],[288,0,288,1],[300,0,300,1],[312,0,312,1],[324,0,324,1],[336,0,336,1],[348,0,348,1],[360,0,360,1],[372,0,372,1],[384,0,384,1],[396,0,396,1],[408,0,408,1],[420,0,420,1],[432,0,432,1],[444,0,444,1],[456,0,456,1],[468,0,468,1],[480,0,480,1],[495,1,496,9],[497,1,497,8],[498,1,498,8],[500,1,500,4],[501,1,501,3],[502,1,502,3],[503,1,503,3],[505,1,505,3]]}
["MARINHA DO BRASIL\n"]
["Noc ANTARES"]
[]
["CALCULO DO OXIGÊNIO DISSOLVIDO NA ÁGUA DO MAR"]
[]
[]
[]
["PERNADA :"]
["PERÍODO :"]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G12]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J12]*112/[.J$499])*([.H12]/273)","v":0},{"$f":"of:=[.K12]/[.M12]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F12]+273.15))+143.3483*LN(([.F12]+273.15)/100)-21.8492*(([.F12]+273.15)/100)+[.I12]*(-0.033096+0.014259*(([.F12]+273.15)/100)-0.0017*((([.F12]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G13]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J13]*112/[.J$499])*([.H13]/273)","v":0},{"$f":"of:=[.K13]/[.M13]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F13]+273.15))+143.3483*LN(([.F13]+273.15)/100)-21.8492*(([.F13]+273.15)/100)+[.I13]*(-0.033096+0.014259*(([.F13]+273.15)/100)-0.0017*((([.F13]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G14]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J14]*112/[.J$499])*([.H14]/273)","v":0},{"$f":"of:=[.K14]/[.M14]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F14]+273.15))+143.3483*LN(([.F14]+273.15)/100)-21.8492*(([.F14]+273.15)/100)+[.I14]*(-0.033096+0.014259*(([.F14]+273.15)/100)-0.0017*((([.F14]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G15]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J15]*112/[.J$499])*([.H15]/273)","v":0},{"$f":"of:=[.K15]/[.M15]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F15]+273.15))+143.3483*LN(([.F15]+273.15)/100)-21.8492*(([.F15]+273.15)/100)+[.I15]*(-0.033096+0.014259*(([.F15]+273.15)/100)-0.0017*((([.F15]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G16]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J16]*112/[.J$499])*([.H16]/273)","v":0},{"$f":"of:=[.K16]/[.M16]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F16]+273.15))+143.3483*LN(([.F16]+273.15)/100)-21.8492*(([.F16]+273.15)/100)+[.I16]*(-0.033096+0.014259*(([.F16]+273.15)/100)-0.0017*((([.F16]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G17]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J17]*112/[.J$499])*([.H17]/273)","v":0},{"$f":"of:=[.K17]/[.M17]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F17]+273.15))+143.3483*LN(([.F17]+273.15)/100)-21.8492*(([.F17]+273.15)/100)+[.I17]*(-0.033096+0.014259*(([.F17]+273.15)/100)-0.0017*((([.F17]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G18]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J18]*112/[.J$499])*([.H18]/273)","v":0},{"$f":"of:=[.K18]/[.M18]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F18]+273.15))+143.3483*LN(([.F18]+273.15)/100)-21.8492*(([.F18]+273.15)/100)+[.I18]*(-0.033096+0.014259*(([.F18]+273.15)/100)-0.0017*((([.F18]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G19]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J19]*112/[.J$499])*([.H19]/273)","v":0},{"$f":"of:=[.K19]/[.M19]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F19]+273.15))+143.3483*LN(([.F19]+273.15)/100)-21.8492*(([.F19]+273.15)/100)+[.I19]*(-0.033096+0.014259*(([.F19]+273.15)/100)-0.0017*((([.F19]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G20]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J20]*112/[.J$499])*([.H20]/273)","v":0},{"$f":"of:=[.K20]/[.M20]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F20]+273.15))+143.3483*LN(([.F20]+273.15)/100)-21.8492*(([.F20]+273.15)/100)+[.I20]*(-0.033096+0.014259*(([.F20]+273.15)/100)-0.0017*((([.F20]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G21]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J21]*112/[.J$499])*([.H21]/273)","v":0},{"$f":"of:=[.K21]/[.M21]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F21]+273.15))+143.3483*LN(([.F21]+273.15)/100)-21.8492*(([.F21]+273.15)/100)+[.I21]*(-0.033096+0.014259*(([.F21]+273.15)/100)-0.0017*((([.F21]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.F24]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J24]*112/[.J$499])*([.H24]/273)","v":0},{"$f":"of:=[.K24]/[.M24]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F24]+273.15))+143.3483*LN(([.F24]+273.15)/100)-21.8492*(([.F24]+273.15)/100)+[.I24]*(-0.033096+0.014259*(([.F24]+273.15)/100)-0.0017*((([.F24]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.F25]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J25]*112/[.J$499])*([.H25]/273)","v":0},{"$f":"of:=[.K25]/[.M25]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F25]+273.15))+143.3483*LN(([.F25]+273.15)/100)-21.8492*(([.F25]+273.15)/100)+[.I25]*(-0.033096+0.014259*(([.F25]+273.15)/100)-0.0017*((([.F25]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G26]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J26]*112/[.J$499])*([.H26]/273)","v":0},{"$f":"of:=[.K26]/[.M26]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F26]+273.15))+143.3483*LN(([.F26]+273.15)/100)-21.8492*(([.F26]+273.15)/100)+[.I26]*(-0.033096+0.014259*(([.F26]+273.15)/100)-0.0017*((([.F26]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G27]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J27]*112/[.J$499])*([.H27]/273)","v":0},{"$f":"of:=[.K27]/[.M27]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F27]+273.15))+143.3483*LN(([.F27]+273.15)/100)-21.8492*(([.F27]+273.15)/100)+[.I27]*(-0.033096+0.014259*(([.F27]+273.15)/100)-0.0017*((([.F27]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G28]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J28]*112/[.J$499])*([.H28]/273)","v":0},{"$f":"of:=[.K28]/[.M28]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F28]+273.15))+143.3483*LN(([.F28]+273.15)/100)-21.8492*(([.F28]+273.15)/100)+[.I28]*(-0.033096+0.014259*(([.F28]+273.15)/100)-0.0017*((([.F28]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G29]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J29]*112/[.J$499])*([.H29]/273)","v":0},{"$f":"of:=[.K29]/[.M29]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F29]+273.15))+143.3483*LN(([.F29]+273.15)/100)-21.8492*(([.F29]+273.15)/100)+[.I29]*(-0.033096+0.014259*(([.F29]+273.15)/100)-0.0017*((([.F29]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G30]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J30]*112/[.J$499])*([.H30]/273)","v":0},{"$f":"of:=[.K30]/[.M30]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F30]+273.15))+143.3483*LN(([.F30]+273.15)/100)-21.8492*(([.F30]+273.15)/100)+[.I30]*(-0.033096+0.014259*(([.F30]+273.15)/100)-0.0017*((([.F30]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G31]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J31]*112/[.J$499])*([.H31]/273)","v":0},{"$f":"of:=[.K31]/[.M31]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F31]+273.15))+143.3483*LN(([.F31]+273.15)/100)-21.8492*(([.F31]+273.15)/100)+[.I31]*(-0.033096+0.014259*(([.F31]+273.15)/100)-0.0017*((([.F31]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G32]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J32]*112/[.J$499])*([.H32]/273)","v":0},{"$f":"of:=[.K32]/[.M32]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F32]+273.15))+143.3483*LN(([.F32]+273.15)/100)-21.8492*(([.F32]+273.15)/100)+[.I32]*(-0.033096+0.014259*(([.F32]+273.15)/100)-0.0017*((([.F32]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G33]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J33]*112/[.J$499])*([.H33]/273)","v":0},{"$f":"of:=[.K33]/[.M33]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F33]+273.15))+143.3483*LN(([.F33]+273.15)/100)-21.8492*(([.F33]+273.15)/100)+[.I33]*(-0.033096+0.014259*(([.F33]+273.15)/100)-0.0017*((([.F33]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G36]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J36]*112/[.J$499])*([.H36]/273)","v":0},{"$f":"of:=[.K36]/[.M36]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F36]+273.15))+143.3483*LN(([.F36]+273.15)/100)-21.8492*(([.F36]+273.15)/100)+[.I36]*(-0.033096+0.014259*(([.F36]+273.15)/100)-0.0017*((([.F36]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G37]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J37]*112/[.J$499])*([.H37]/273)","v":0},{"$f":"of:=[.K37]/[.M37]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F37]+273.15))+143.3483*LN(([.F37]+273.15)/100)-21.8492*(([.F37]+273.15)/100)+[.I37]*(-0.033096+0.014259*(([.F37]+273.15)/100)-0.0017*((([.F37]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G38]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J38]*112/[.J$499])*([.H38]/273)","v":0},{"$f":"of:=[.K38]/[.M38]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F38]+273.15))+143.3483*LN(([.F38]+273.15)/100)-21.8492*(([.F38]+273.15)/100)+[.I38]*(-0.033096+0.014259*(([.F38]+273.15)/100)-0.0017*((([.F38]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G39]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J39]*112/[.J$499])*([.H39]/273)","v":0},{"$f":"of:=[.K39]/[.M39]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F39]+273.15))+143.3483*LN(([.F39]+273.15)/100)-21.8492*(([.F39]+273.15)/100)+[.I39]*(-0.033096+0.014259*(([.F39]+273.15)/100)-0.0017*((([.F39]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G40]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J40]*112/[.J$499])*([.H40]/273)","v":0},{"$f":"of:=[.K40]/[.M40]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F40]+273.15))+143.3483*LN(([.F40]+273.15)/100)-21.8492*(([.F40]+273.15)/100)+[.I40]*(-0.033096+0.014259*(([.F40]+273.15)/100)-0.0017*((([.F40]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G41]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J41]*112/[.J$499])*([.H41]/273)","v":0},{"$f":"of:=[.K41]/[.M41]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F41]+273.15))+143.3483*LN(([.F41]+273.15)/100)-21.8492*(([.F41]+273.15)/100)+[.I41]*(-0.033096+0.014259*(([.F41]+273.15)/100)-0.0017*((([.F41]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G42]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J42]*112/[.J$499])*([.H42]/273)","v":0},{"$f":"of:=[.K42]/[.M42]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F42]+273.15))+143.3483*LN(([.F42]+273.15)/100)-21.8492*(([.F42]+273.15)/100)+[.I42]*(-0.033096+0.014259*(([.F42]+273.15)/100)-0.0017*((([.F42]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G43]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J43]*112/[.J$499])*([.H43]/273)","v":0},{"$f":"of:=[.K43]/[.M43]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F43]+273.15))+143.3483*LN(([.F43]+273.15)/100)-21.8492*(([.F43]+273.15)/100)+[.I43]*(-0.033096+0.014259*(([.F43]+273.15)/100)-0.0017*((([.F43]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G44]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J44]*112/[.J$499])*([.H44]/273)","v":0},{"$f":"of:=[.K44]/[.M44]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F44]+273.15))+143.3483*LN(([.F44]+273.15)/100)-21.8492*(([.F44]+273.15)/100)+[.I44]*(-0.033096+0.014259*(([.F44]+273.15)/100)-0.0017*((([.F44]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G45]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J45]*112/[.J$499])*([.H45]/273)","v":0},{"$f":"of:=[.K45]/[.M45]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F45]+273.15))+143.3483*LN(([.F45]+273.15)/100)-21.8492*(([.F45]+273.15)/100)+[.I45]*(-0.033096+0.014259*(([.F45]+273.15)/100)-0.0017*((([.F45]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G48]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J48]*112/[.J$499])*([.H48]/273)","v":0},{"$f":"of:=[.K48]/[.M48]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F48]+273.15))+143.3483*LN(([.F48]+273.15)/100)-21.8492*(([.F48]+273.15)/100)+[.I48]*(-0.033096+0.014259*(([.F48]+273.15)/100)-0.0017*((([.F48]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G49]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J49]*112/[.J$499])*([.H49]/273)","v":0},{"$f":"of:=[.K49]/[.M49]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F49]+273.15))+143.3483*LN(([.F49]+273.15)/100)-21.8492*(([.F49]+273.15)/100)+[.I49]*(-0.033096+0.014259*(([.F49]+273.15)/100)-0.0017*((([.F49]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G50]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J50]*112/[.J$499])*([.H50]/273)","v":0},{"$f":"of:=[.K50]/[.M50]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F50]+273.15))+143.3483*LN(([.F50]+273.15)/100)-21.8492*(([.F50]+273.15)/100)+[.I50]*(-0.033096+0.014259*(([.F50]+273.15)/100)-0.0017*((([.F50]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G51]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J51]*112/[.J$499])*([.H51]/273)","v":0},{"$f":"of:=[.K51]/[.M51]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F51]+273.15))+143.3483*LN(([.F51]+273.15)/100)-21.8492*(([.F51]+273.15)/100)+[.I51]*(-0.033096+0.014259*(([.F51]+273.15)/100)-0.0017*((([.F51]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G52]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J52]*112/[.J$499])*([.H52]/273)","v":0},{"$f":"of:=[.K52]/[.M52]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F52]+273.15))+143.3483*LN(([.F52]+273.15)/100)-21.8492*(([.F52]+273.15)/100)+[.I52]*(-0.033096+0.014259*(([.F52]+273.15)/100)-0.0017*((([.F52]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G53]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J53]*112/[.J$499])*([.H53]/273)","v":0},{"$f":"of:=[.K53]/[.M53]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F53]+273.15))+143.3483*LN(([.F53]+273.15)/100)-21.8492*(([.F53]+273.15)/100)+[.I53]*(-0.033096+0.014259*(([.F53]+273.15)/100)-0.0017*((([.F53]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G54]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J54]*112/[.J$499])*([.H54]/273)","v":0},{"$f":"of:=[.K54]/[.M54]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F54]+273.15))+143.3483*LN(([.F54]+273.15)/100)-21.8492*(([.F54]+273.15)/100)+[.I54]*(-0.033096+0.014259*(([.F54]+273.15)/100)-0.0017*((([.F54]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G55]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J55]*112/[.J$499])*([.H55]/273)","v":0},{"$f":"of:=[.K55]/[.M55]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F55]+273.15))+143.3483*LN(([.F55]+273.15)/100)-21.8492*(([.F55]+273.15)/100)+[.I55]*(-0.033096+0.014259*(([.F55]+273.15)/100)-0.0017*((([.F55]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G56]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J56]*112/[.J$499])*([.H56]/273)","v":0},{"$f":"of:=[.K56]/[.M56]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F56]+273.15))+143.3483*LN(([.F56]+273.15)/100)-21.8492*(([.F56]+273.15)/100)+[.I56]*(-0.033096+0.014259*(([.F56]+273.15)/100)-0.0017*((([.F56]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G57]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J57]*112/[.J$499])*([.H57]/273)","v":0},{"$f":"of:=[.K57]/[.M57]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F57]+273.15))+143.3483*LN(([.F57]+273.15)/100)-21.8492*(([.F57]+273.15)/100)+[.I57]*(-0.033096+0.014259*(([.F57]+273.15)/100)-0.0017*((([.F57]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G60]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J60]*112/[.J$499])*([.H60]/273)","v":0},{"$f":"of:=[.K60]/[.M60]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F60]+273.15))+143.3483*LN(([.F60]+273.15)/100)-21.8492*(([.F60]+273.15)/100)+[.I60]*(-0.033096+0.014259*(([.F60]+273.15)/100)-0.0017*((([.F60]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G61]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J61]*112/[.J$499])*([.H61]/273)","v":0},{"$f":"of:=[.K61]/[.M61]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F61]+273.15))+143.3483*LN(([.F61]+273.15)/100)-21.8492*(([.F61]+273.15)/100)+[.I61]*(-0.033096+0.014259*(([.F61]+273.15)/100)-0.0017*((([.F61]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G62]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J62]*112/[.J$499])*([.H62]/273)","v":0},{"$f":"of:=[.K62]/[.M62]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F62]+273.15))+143.3483*LN(([.F62]+273.15)/100)-21.8492*(([.F62]+273.15)/100)+[.I62]*(-0.033096+0.014259*(([.F62]+273.15)/100)-0.0017*((([.F62]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G63]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J63]*112/[.J$499])*([.H63]/273)","v":0},{"$f":"of:=[.K63]/[.M63]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F63]+273.15))+143.3483*LN(([.F63]+273.15)/100)-21.8492*(([.F63]+273.15)/100)+[.I63]*(-0.033096+0.014259*(([.F63]+273.15)/100)-0.0017*((([.F63]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G64]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J64]*112/[.J$499])*([.H64]/273)","v":0},{"$f":"of:=[.K64]/[.M64]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F64]+273.15))+143.3483*LN(([.F64]+273.15)/100)-21.8492*(([.F64]+273.15)/100)+[.I64]*(-0.033096+0.014259*(([.F64]+273.15)/100)-0.0017*((([.F64]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G65]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J65]*112/[.J$499])*([.H65]/273)","v":0},{"$f":"of:=[.K65]/[.M65]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F65]+273.15))+143.3483*LN(([.F65]+273.15)/100)-21.8492*(([.F65]+273.15)/100)+[.I65]*(-0.033096+0.014259*(([.F65]+273.15)/100)-0.0017*((([.F65]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G66]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J66]*112/[.J$499])*([.H66]/273)","v":0},{"$f":"of:=[.K66]/[.M66]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F66]+273.15))+143.3483*LN(([.F66]+273.15)/100)-21.8492*(([.F66]+273.15)/100)+[.I66]*(-0.033096+0.014259*(([.F66]+273.15)/100)-0.0017*((([.F66]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G67]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J67]*112/[.J$499])*([.H67]/273)","v":0},{"$f":"of:=[.K67]/[.M67]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F67]+273.15))+143.3483*LN(([.F67]+273.15)/100)-21.8492*(([.F67]+273.15)/100)+[.I67]*(-0.033096+0.014259*(([.F67]+273.15)/100)-0.0017*((([.F67]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G68]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J68]*112/[.J$499])*([.H68]/273)","v":0},{"$f":"of:=[.K68]/[.M68]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F68]+273.15))+143.3483*LN(([.F68]+273.15)/100)-21.8492*(([.F68]+273.15)/100)+[.I68]*(-0.033096+0.014259*(([.F68]+273.15)/100)-0.0017*((([.F68]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G69]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J69]*112/[.J$499])*([.H69]/273)","v":0},{"$f":"of:=[.K69]/[.M69]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F69]+273.15))+143.3483*LN(([.F69]+273.15)/100)-21.8492*(([.F69]+273.15)/100)+[.I69]*(-0.033096+0.014259*(([.F69]+273.15)/100)-0.0017*((([.F69]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G72]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J72]*112/[.J$499])*([.H72]/273)","v":0},{"$f":"of:=[.K72]/[.M72]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F72]+273.15))+143.3483*LN(([.F72]+273.15)/100)-21.8492*(([.F72]+273.15)/100)+[.I72]*(-0.033096+0.014259*(([.F72]+273.15)/100)-0.0017*((([.F72]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G73]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J73]*112/[.J$499])*([.H73]/273)","v":0},{"$f":"of:=[.K73]/[.M73]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F73]+273.15))+143.3483*LN(([.F73]+273.15)/100)-21.8492*(([.F73]+273.15)/100)+[.I73]*(-0.033096+0.014259*(([.F73]+273.15)/100)-0.0017*((([.F73]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G74]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J74]*112/[.J$499])*([.H74]/273)","v":0},{"$f":"of:=[.K74]/[.M74]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F74]+273.15))+143.3483*LN(([.F74]+273.15)/100)-21.8492*(([.F74]+273.15)/100)+[.I74]*(-0.033096+0.014259*(([.F74]+273.15)/100)-0.0017*((([.F74]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G75]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J75]*112/[.J$499])*([.H75]/273)","v":0},{"$f":"of:=[.K75]/[.M75]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F75]+273.15))+143.3483*LN(([.F75]+273.15)/100)-21.8492*(([.F75]+273.15)/100)+[.I75]*(-0.033096+0.014259*(([.F75]+273.15)/100)-0.0017*((([.F75]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G76]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J76]*112/[.J$499])*([.H76]/273)","v":0},{"$f":"of:=[.K76]/[.M76]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F76]+273.15))+143.3483*LN(([.F76]+273.15)/100)-21.8492*(([.F76]+273.15)/100)+[.I76]*(-0.033096+0.014259*(([.F76]+273.15)/100)-0.0017*((([.F76]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G77]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J77]*112/[.J$499])*([.H77]/273)","v":0},{"$f":"of:=[.K77]/[.M77]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F77]+273.15))+143.3483*LN(([.F77]+273.15)/100)-21.8492*(([.F77]+273.15)/100)+[.I77]*(-0.033096+0.014259*(([.F77]+273.15)/100)-0.0017*((([.F77]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G78]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J78]*112/[.J$499])*([.H78]/273)","v":0},{"$f":"of:=[.K78]/[.M78]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F78]+273.15))+143.3483*LN(([.F78]+273.15)/100)-21.8492*(([.F78]+273.15)/100)+[.I78]*(-0.033096+0.014259*(([.F78]+273.15)/100)-0.0017*((([.F78]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G79]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J79]*112/[.J$499])*([.H79]/273)","v":0},{"$f":"of:=[.K79]/[.M79]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F79]+273.15))+143.3483*LN(([.F79]+273.15)/100)-21.8492*(([.F79]+273.15)/100)+[.I79]*(-0.033096+0.014259*(([.F79]+273.15)/100)-0.0017*((([.F79]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G80]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J80]*112/[.J$499])*([.H80]/273)","v":0},{"$f":"of:=[.K80]/[.M80]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F80]+273.15))+143.3483*LN(([.F80]+273.15)/100)-21.8492*(([.F80]+273.15)/100)+[.I80]*(-0.033096+0.014259*(([.F80]+273.15)/100)-0.0017*((([.F80]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G81]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J81]*112/[.J$499])*([.H81]/273)","v":0},{"$f":"of:=[.K81]/[.M81]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F81]+273.15))+143.3483*LN(([.F81]+273.15)/100)-21.8492*(([.F81]+273.15)/100)+[.I81]*(-0.033096+0.014259*(([.F81]+273.15)/100)-0.0017*((([.F81]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G84]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J84]*112/[.J$499])*([.H84]/273)","v":0},{"$f":"of:=[.K84]/[.M84]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F84]+273.15))+143.3483*LN(([.F84]+273.15)/100)-21.8492*(([.F84]+273.15)/100)+[.I84]*(-0.033096+0.014259*(([.F84]+273.15)/100)-0.0017*((([.F84]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G85]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J85]*112/[.J$499])*([.H85]/273)","v":0},{"$f":"of:=[.K85]/[.M85]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F85]+273.15))+143.3483*LN(([.F85]+273.15)/100)-21.8492*(([.F85]+273.15)/100)+[.I85]*(-0.033096+0.014259*(([.F85]+273.15)/100)-0.0017*((([.F85]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G86]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J86]*112/[.J$499])*([.H86]/273)","v":0},{"$f":"of:=[.K86]/[.M86]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F86]+273.15))+143.3483*LN(([.F86]+273.15)/100)-21.8492*(([.F86]+273.15)/100)+[.I86]*(-0.033096+0.014259*(([.F86]+273.15)/100)-0.0017*((([.F86]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G87]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J87]*112/[.J$499])*([.H87]/273)","v":0},{"$f":"of:=[.K87]/[.M87]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F87]+273.15))+143.3483*LN(([.F87]+273.15)/100)-21.8492*(([.F87]+273.15)/100)+[.I87]*(-0.033096+0.014259*(([.F87]+273.15)/100)-0.0017*((([.F87]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G88]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J88]*112/[.J$499])*([.H88]/273)","v":0},{"$f":"of:=[.K88]/[.M88]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F88]+273.15))+143.3483*LN(([.F88]+273.15)/100)-21.8492*(([.F88]+273.15)/100)+[.I88]*(-0.033096+0.014259*(([.F88]+273.15)/100)-0.0017*((([.F88]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G89]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J89]*112/[.J$499])*([.H89]/273)","v":0},{"$f":"of:=[.K89]/[.M89]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F89]+273.15))+143.3483*LN(([.F89]+273.15)/100)-21.8492*(([.F89]+273.15)/100)+[.I89]*(-0.033096+0.014259*(([.F89]+273.15)/100)-0.0017*((([.F89]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G90]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J90]*112/[.J$499])*([.H90]/273)","v":0},{"$f":"of:=[.K90]/[.M90]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F90]+273.15))+143.3483*LN(([.F90]+273.15)/100)-21.8492*(([.F90]+273.15)/100)+[.I90]*(-0.033096+0.014259*(([.F90]+273.15)/100)-0.0017*((([.F90]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G91]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J91]*112/[.J$499])*([.H91]/273)","v":0},{"$f":"of:=[.K91]/[.M91]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F91]+273.15))+143.3483*LN(([.F91]+273.15)/100)-21.8492*(([.F91]+273.15)/100)+[.I91]*(-0.033096+0.014259*(([.F91]+273.15)/100)-0.0017*((([.F91]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G92]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J92]*112/[.J$499])*([.H92]/273)","v":0},{"$f":"of:=[.K92]/[.M92]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F92]+273.15))+143.3483*LN(([.F92]+273.15)/100)-21.8492*(([.F92]+273.15)/100)+[.I92]*(-0.033096+0.014259*(([.F92]+273.15)/100)-0.0017*((([.F92]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G93]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J93]*112/[.J$499])*([.H93]/273)","v":0},{"$f":"of:=[.K93]/[.M93]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F93]+273.15))+143.3483*LN(([.F93]+273.15)/100)-21.8492*(([.F93]+273.15)/100)+[.I93]*(-0.033096+0.014259*(([.F93]+273.15)/100)-0.0017*((([.F93]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G96]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J96]*112/[.J$499])*([.H96]/273)","v":0},{"$f":"of:=[.K96]/[.M96]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F96]+273.15))+143.3483*LN(([.F96]+273.15)/100)-21.8492*(([.F96]+273.15)/100)+[.I96]*(-0.033096+0.014259*(([.F96]+273.15)/100)-0.0017*((([.F96]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G97]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J97]*112/[.J$499])*([.H97]/273)","v":0},{"$f":"of:=[.K97]/[.M97]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F97]+273.15))+143.3483*LN(([.F97]+273.15)/100)-21.8492*(([.F97]+273.15)/100)+[.I97]*(-0.033096+0.014259*(([.F97]+273.15)/100)-0.0017*((([.F97]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G98]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J98]*112/[.J$499])*([.H98]/273)","v":0},{"$f":"of:=[.K98]/[.M98]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F98]+273.15))+143.3483*LN(([.F98]+273.15)/100)-21.8492*(([.F98]+273.15)/100)+[.I98]*(-0.033096+0.014259*(([.F98]+273.15)/100)-0.0017*((([.F98]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G99]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J99]*112/[.J$499])*([.H99]/273)","v":0},{"$f":"of:=[.K99]/[.M99]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F99]+273.15))+143.3483*LN(([.F99]+273.15)/100)-21.8492*(([.F99]+273.15)/100)+[.I99]*(-0.033096+0.014259*(([.F99]+273.15)/100)-0.0017*((([.F99]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G100]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J100]*112/[.J$499])*([.H100]/273)","v":0},{"$f":"of:=[.K100]/[.M100]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F100]+273.15))+143.3483*LN(([.F100]+273.15)/100)-21.8492*(([.F100]+273.15)/100)+[.I100]*(-0.033096+0.014259*(([.F100]+273.15)/100)-0.0017*((([.F100]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G101]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J101]*112/[.J$499])*([.H101]/273)","v":0},{"$f":"of:=[.K101]/[.M101]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F101]+273.15))+143.3483*LN(([.F101]+273.15)/100)-21.8492*(([.F101]+273.15)/100)+[.I101]*(-0.033096+0.014259*(([.F101]+273.15)/100)-0.0017*((([.F101]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G102]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J102]*112/[.J$499])*([.H102]/273)","v":0},{"$f":"of:=[.K102]/[.M102]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F102]+273.15))+143.3483*LN(([.F102]+273.15)/100)-21.8492*(([.F102]+273.15)/100)+[.I102]*(-0.033096+0.014259*(([.F102]+273.15)/100)-0.0017*((([.F102]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G103]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J103]*112/[.J$499])*([.H103]/273)","v":0},{"$f":"of:=[.K103]/[.M103]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F103]+273.15))+143.3483*LN(([.F103]+273.15)/100)-21.8492*(([.F103]+273.15)/100)+[.I103]*(-0.033096+0.014259*(([.F103]+273.15)/100)-0.0017*((([.F103]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G104]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J104]*112/[.J$499])*([.H104]/273)","v":0},{"$f":"of:=[.K104]/[.M104]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F104]+273.15))+143.3483*LN(([.F104]+273.15)/100)-21.8492*(([.F104]+273.15)/100)+[.I104]*(-0.033096+0.014259*(([.F104]+273.15)/100)-0.0017*((([.F104]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G105]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J105]*112/[.J$499])*([.H105]/273)","v":0},{"$f":"of:=[.K105]/[.M105]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F105]+273.15))+143.3483*LN(([.F105]+273.15)/100)-21.8492*(([.F105]+273.15)/100)+[.I105]*(-0.033096+0.014259*(([.F105]+273.15)/100)-0.0017*((([.F105]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G108]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J108]*112/[.J$499])*([.H108]/273)","v":0},{"$f":"of:=[.K108]/[.M108]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F108]+273.15))+143.3483*LN(([.F108]+273.15)/100)-21.8492*(([.F108]+273.15)/100)+[.I108]*(-0.033096+0.014259*(([.F108]+273.15)/100)-0.0017*((([.F108]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G109]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J109]*112/[.J$499])*([.H109]/273)","v":0},{"$f":"of:=[.K109]/[.M109]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F109]+273.15))+143.3483*LN(([.F109]+273.15)/100)-21.8492*(([.F109]+273.15)/100)+[.I109]*(-0.033096+0.014259*(([.F109]+273.15)/100)-0.0017*((([.F109]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G110]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J110]*112/[.J$499])*([.H110]/273)","v":0},{"$f":"of:=[.K110]/[.M110]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F110]+273.15))+143.3483*LN(([.F110]+273.15)/100)-21.8492*(([.F110]+273.15)/100)+[.I110]*(-0.033096+0.014259*(([.F110]+273.15)/100)-0.0017*((([.F110]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G111]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J111]*112/[.J$499])*([.H111]/273)","v":0},{"$f":"of:=[.K111]/[.M111]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F111]+273.15))+143.3483*LN(([.F111]+273.15)/100)-21.8492*(([.F111]+273.15)/100)+[.I111]*(-0.033096+0.014259*(([.F111]+273.15)/100)-0.0017*((([.F111]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G112]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J112]*112/[.J$499])*([.H112]/273)","v":0},{"$f":"of:=[.K112]/[.M112]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F112]+273.15))+143.3483*LN(([.F112]+273.15)/100)-21.8492*(([.F112]+273.15)/100)+[.I112]*(-0.033096+0.014259*(([.F112]+273.15)/100)-0.0017*((([.F112]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G113]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J113]*112/[.J$499])*([.H113]/273)","v":0},{"$f":"of:=[.K113]/[.M113]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F113]+273.15))+143.3483*LN(([.F113]+273.15)/100)-21.8492*(([.F113]+273.15)/100)+[.I113]*(-0.033096+0.014259*(([.F113]+273.15)/100)-0.0017*((([.F113]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G114]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J114]*112/[.J$499])*([.H114]/273)","v":0},{"$f":"of:=[.K114]/[.M114]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F114]+273.15))+143.3483*LN(([.F114]+273.15)/100)-21.8492*(([.F114]+273.15)/100)+[.I114]*(-0.033096+0.014259*(([.F114]+273.15)/100)-0.0017*((([.F114]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G115]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J115]*112/[.J$499])*([.H115]/273)","v":0},{"$f":"of:=[.K115]/[.M115]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F115]+273.15))+143.3483*LN(([.F115]+273.15)/100)-21.8492*(([.F115]+273.15)/100)+[.I115]*(-0.033096+0.014259*(([.F115]+273.15)/100)-0.0017*((([.F115]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G116]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J116]*112/[.J$499])*([.H116]/273)","v":0},{"$f":"of:=[.K116]/[.M116]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F116]+273.15))+143.3483*LN(([.F116]+273.15)/100)-21.8492*(([.F116]+273.15)/100)+[.I116]*(-0.033096+0.014259*(([.F116]+273.15)/100)-0.0017*((([.F116]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G117]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J117]*112/[.J$499])*([.H117]/273)","v":0},{"$f":"of:=[.K117]/[.M117]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F117]+273.15))+143.3483*LN(([.F117]+273.15)/100)-21.8492*(([.F117]+273.15)/100)+[.I117]*(-0.033096+0.014259*(([.F117]+273.15)/100)-0.0017*((([.F117]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G120]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J120]*112/[.J$499])*([.H120]/273)","v":0},{"$f":"of:=[.K120]/[.M120]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F120]+273.15))+143.3483*LN(([.F120]+273.15)/100)-21.8492*(([.F120]+273.15)/100)+[.I120]*(-0.033096+0.014259*(([.F120]+273.15)/100)-0.0017*((([.F120]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G121]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J121]*112/[.J$499])*([.H121]/273)","v":0},{"$f":"of:=[.K121]/[.M121]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F121]+273.15))+143.3483*LN(([.F121]+273.15)/100)-21.8492*(([.F121]+273.15)/100)+[.I121]*(-0.033096+0.014259*(([.F121]+273.15)/100)-0.0017*((([.F121]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G122]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J122]*112/[.J$499])*([.H122]/273)","v":0},{"$f":"of:=[.K122]/[.M122]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F122]+273.15))+143.3483*LN(([.F122]+273.15)/100)-21.8492*(([.F122]+273.15)/100)+[.I122]*(-0.033096+0.014259*(([.F122]+273.15)/100)-0.0017*((([.F122]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G123]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J123]*112/[.J$499])*([.H123]/273)","v":0},{"$f":"of:=[.K123]/[.M123]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F123]+273.15))+143.3483*LN(([.F123]+273.15)/100)-21.8492*(([.F123]+273.15)/100)+[.I123]*(-0.033096+0.014259*(([.F123]+273.15)/100)-0.0017*((([.F123]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G124]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J124]*112/[.J$499])*([.H124]/273)","v":0},{"$f":"of:=[.K124]/[.M124]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F124]+273.15))+143.3483*LN(([.F124]+273.15)/100)-21.8492*(([.F124]+273.15)/100)+[.I124]*(-0.033096+0.014259*(([.F124]+273.15)/100)-0.0017*((([.F124]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G125]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J125]*112/[.J$499])*([.H125]/273)","v":0},{"$f":"of:=[.K125]/[.M125]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F125]+273.15))+143.3483*LN(([.F125]+273.15)/100)-21.8492*(([.F125]+273.15)/100)+[.I125]*(-0.033096+0.014259*(([.F125]+273.15)/100)-0.0017*((([.F125]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G126]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J126]*112/[.J$499])*([.H126]/273)","v":0},{"$f":"of:=[.K126]/[.M126]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F126]+273.15))+143.3483*LN(([.F126]+273.15)/100)-21.8492*(([.F126]+273.15)/100)+[.I126]*(-0.033096+0.014259*(([.F126]+273.15)/100)-0.0017*((([.F126]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G127]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J127]*112/[.J$499])*([.H127]/273)","v":0},{"$f":"of:=[.K127]/[.M127]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F127]+273.15))+143.3483*LN(([.F127]+273.15)/100)-21.8492*(([.F127]+273.15)/100)+[.I127]*(-0.033096+0.014259*(([.F127]+273.15)/100)-0.0017*((([.F127]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G128]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J128]*112/[.J$499])*([.H128]/273)","v":0},{"$f":"of:=[.K128]/[.M128]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F128]+273.15))+143.3483*LN(([.F128]+273.15)/100)-21.8492*(([.F128]+273.15)/100)+[.I128]*(-0.033096+0.014259*(([.F128]+273.15)/100)-0.0017*((([.F128]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G129]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J129]*112/[.J$499])*([.H129]/273)","v":0},{"$f":"of:=[.K129]/[.M129]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F129]+273.15))+143.3483*LN(([.F129]+273.15)/100)-21.8492*(([.F129]+273.15)/100)+[.I129]*(-0.033096+0.014259*(([.F129]+273.15)/100)-0.0017*((([.F129]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G132]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J132]*112/[.J$499])*([.H132]/273)","v":0},{"$f":"of:=[.K132]/[.M132]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F132]+273.15))+143.3483*LN(([.F132]+273.15)/100)-21.8492*(([.F132]+273.15)/100)+[.I132]*(-0.033096+0.014259*(([.F132]+273.15)/100)-0.0017*((([.F132]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G133]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J133]*112/[.J$499])*([.H133]/273)","v":0},{"$f":"of:=[.K133]/[.M133]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F133]+273.15))+143.3483*LN(([.F133]+273.15)/100)-21.8492*(([.F133]+273.15)/100)+[.I133]*(-0.033096+0.014259*(([.F133]+273.15)/100)-0.0017*((([.F133]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G134]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J134]*112/[.J$499])*([.H134]/273)","v":0},{"$f":"of:=[.K134]/[.M134]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F134]+273.15))+143.3483*LN(([.F134]+273.15)/100)-21.8492*(([.F134]+273.15)/100)+[.I134]*(-0.033096+0.014259*(([.F134]+273.15)/100)-0.0017*((([.F134]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G135]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J135]*112/[.J$499])*([.H135]/273)","v":0},{"$f":"of:=[.K135]/[.M135]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F135]+273.15))+143.3483*LN(([.F135]+273.15)/100)-21.8492*(([.F135]+273.15)/100)+[.I135]*(-0.033096+0.014259*(([.F135]+273.15)/100)-0.0017*((([.F135]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G136]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J136]*112/[.J$499])*([.H136]/273)","v":0},{"$f":"of:=[.K136]/[.M136]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F136]+273.15))+143.3483*LN(([.F136]+273.15)/100)-21.8492*(([.F136]+273.15)/100)+[.I136]*(-0.033096+0.014259*(([.F136]+273.15)/100)-0.0017*((([.F136]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G137]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J137]*112/[.J$499])*([.H137]/273)","v":0},{"$f":"of:=[.K137]/[.M137]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F137]+273.15))+143.3483*LN(([.F137]+273.15)/100)-21.8492*(([.F137]+273.15)/100)+[.I137]*(-0.033096+0.014259*(([.F137]+273.15)/100)-0.0017*((([.F137]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G138]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J138]*112/[.J$499])*([.H138]/273)","v":0},{"$f":"of:=[.K138]/[.M138]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F138]+273.15))+143.3483*LN(([.F138]+273.15)/100)-21.8492*(([.F138]+273.15)/100)+[.I138]*(-0.033096+0.014259*(([.F138]+273.15)/100)-0.0017*((([.F138]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G139]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J139]*112/[.J$499])*([.H139]/273)","v":0},{"$f":"of:=[.K139]/[.M139]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F139]+273.15))+143.3483*LN(([.F139]+273.15)/100)-21.8492*(([.F139]+273.15)/100)+[.I139]*(-0.033096+0.014259*(([.F139]+273.15)/100)-0.0017*((([.F139]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G140]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J140]*112/[.J$499])*([.H140]/273)","v":0},{"$f":"of:=[.K140]/[.M140]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F140]+273.15))+143.3483*LN(([.F140]+273.15)/100)-21.8492*(([.F140]+273.15)/100)+[.I140]*(-0.033096+0.014259*(([.F140]+273.15)/100)-0.0017*((([.F140]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G141]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J141]*112/[.J$499])*([.H141]/273)","v":0},{"$f":"of:=[.K141]/[.M141]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F141]+273.15))+143.3483*LN(([.F141]+273.15)/100)-21.8492*(([.F141]+273.15)/100)+[.I141]*(-0.033096+0.014259*(([.F141]+273.15)/100)-0.0017*((([.F141]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G144]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J144]*112/[.J$499])*([.H144]/273)","v":0},{"$f":"of:=[.K144]/[.M144]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F144]+273.15))+143.3483*LN(([.F144]+273.15)/100)-21.8492*(([.F144]+273.15)/100)+[.I144]*(-0.033096+0.014259*(([.F144]+273.15)/100)-0.0017*((([.F144]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G145]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J145]*112/[.J$499])*([.H145]/273)","v":0},{"$f":"of:=[.K145]/[.M145]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F145]+273.15))+143.3483*LN(([.F145]+273.15)/100)-21.8492*(([.F145]+273.15)/100)+[.I145]*(-0.033096+0.014259*(([.F145]+273.15)/100)-0.0017*((([.F145]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G146]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J146]*112/[.J$499])*([.H146]/273)","v":0},{"$f":"of:=[.K146]/[.M146]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F146]+273.15))+143.3483*LN(([.F146]+273.15)/100)-21.8492*(([.F146]+273.15)/100)+[.I146]*(-0.033096+0.014259*(([.F146]+273.15)/100)-0.0017*((([.F146]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G147]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J147]*112/[.J$499])*([.H147]/273)","v":0},{"$f":"of:=[.K147]/[.M147]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F147]+273.15))+143.3483*LN(([.F147]+273.15)/100)-21.8492*(([.F147]+273.15)/100)+[.I147]*(-0.033096+0.014259*(([.F147]+273.15)/100)-0.0017*((([.F147]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G148]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J148]*112/[.J$499])*([.H148]/273)","v":0},{"$f":"of:=[.K148]/[.M148]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F148]+273.15))+143.3483*LN(([.F148]+273.15)/100)-21.8492*(([.F148]+273.15)/100)+[.I148]*(-0.033096+0.014259*(([.F148]+273.15)/100)-0.0017*((([.F148]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G149]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J149]*112/[.J$499])*([.H149]/273)","v":0},{"$f":"of:=[.K149]/[.M149]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F149]+273.15))+143.3483*LN(([.F149]+273.15)/100)-21.8492*(([.F149]+273.15)/100)+[.I149]*(-0.033096+0.014259*(([.F149]+273.15)/100)-0.0017*((([.F149]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G150]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J150]*112/[.J$499])*([.H150]/273)","v":0},{"$f":"of:=[.K150]/[.M150]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F150]+273.15))+143.3483*LN(([.F150]+273.15)/100)-21.8492*(([.F150]+273.15)/100)+[.I150]*(-0.033096+0.014259*(([.F150]+273.15)/100)-0.0017*((([.F150]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G151]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J151]*112/[.J$499])*([.H151]/273)","v":0},{"$f":"of:=[.K151]/[.M151]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F151]+273.15))+143.3483*LN(([.F151]+273.15)/100)-21.8492*(([.F151]+273.15)/100)+[.I151]*(-0.033096+0.014259*(([.F151]+273.15)/100)-0.0017*((([.F151]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G152]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J152]*112/[.J$499])*([.H152]/273)","v":0},{"$f":"of:=[.K152]/[.M152]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F152]+273.15))+143.3483*LN(([.F152]+273.15)/100)-21.8492*(([.F152]+273.15)/100)+[.I152]*(-0.033096+0.014259*(([.F152]+273.15)/100)-0.0017*((([.F152]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G153]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J153]*112/[.J$499])*([.H153]/273)","v":0},{"$f":"of:=[.K153]/[.M153]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F153]+273.15))+143.3483*LN(([.F153]+273.15)/100)-21.8492*(([.F153]+273.15)/100)+[.I153]*(-0.033096+0.014259*(([.F153]+273.15)/100)-0.0017*((([.F153]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G156]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J156]*112/[.J$499])*([.H156]/273)","v":0},{"$f":"of:=[.K156]/[.M156]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F156]+273.15))+143.3483*LN(([.F156]+273.15)/100)-21.8492*(([.F156]+273.15)/100)+[.I156]*(-0.033096+0.014259*(([.F156]+273.15)/100)-0.0017*((([.F156]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G157]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J157]*112/[.J$499])*([.H157]/273)","v":0},{"$f":"of:=[.K157]/[.M157]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F157]+273.15))+143.3483*LN(([.F157]+273.15)/100)-21.8492*(([.F157]+273.15)/100)+[.I157]*(-0.033096+0.014259*(([.F157]+273.15)/100)-0.0017*((([.F157]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G158]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J158]*112/[.J$499])*([.H158]/273)","v":0},{"$f":"of:=[.K158]/[.M158]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F158]+273.15))+143.3483*LN(([.F158]+273.15)/100)-21.8492*(([.F158]+273.15)/100)+[.I158]*(-0.033096+0.014259*(([.F158]+273.15)/100)-0.0017*((([.F158]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G159]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J159]*112/[.J$499])*([.H159]/273)","v":0},{"$f":"of:=[.K159]/[.M159]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F159]+273.15))+143.3483*LN(([.F159]+273.15)/100)-21.8492*(([.F159]+273.15)/100)+[.I159]*(-0.033096+0.014259*(([.F159]+273.15)/100)-0.0017*((([.F159]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G160]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J160]*112/[.J$499])*([.H160]/273)","v":0},{"$f":"of:=[.K160]/[.M160]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F160]+273.15))+143.3483*LN(([.F160]+273.15)/100)-21.8492*(([.F160]+273.15)/100)+[.I160]*(-0.033096+0.014259*(([.F160]+273.15)/100)-0.0017*((([.F160]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G161]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J161]*112/[.J$499])*([.H161]/273)","v":0},{"$f":"of:=[.K161]/[.M161]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F161]+273.15))+143.3483*LN(([.F161]+273.15)/100)-21.8492*(([.F161]+273.15)/100)+[.I161]*(-0.033096+0.014259*(([.F161]+273.15)/100)-0.0017*((([.F161]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G162]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J162]*112/[.J$499])*([.H162]/273)","v":0},{"$f":"of:=[.K162]/[.M162]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F162]+273.15))+143.3483*LN(([.F162]+273.15)/100)-21.8492*(([.F162]+273.15)/100)+[.I162]*(-0.033096+0.014259*(([.F162]+273.15)/100)-0.0017*((([.F162]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G163]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J163]*112/[.J$499])*([.H163]/273)","v":0},{"$f":"of:=[.K163]/[.M163]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F163]+273.15))+143.3483*LN(([.F163]+273.15)/100)-21.8492*(([.F163]+273.15)/100)+[.I163]*(-0.033096+0.014259*(([.F163]+273.15)/100)-0.0017*((([.F163]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G164]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J164]*112/[.J$499])*([.H164]/273)","v":0},{"$f":"of:=[.K164]/[.M164]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F164]+273.15))+143.3483*LN(([.F164]+273.15)/100)-21.8492*(([.F164]+273.15)/100)+[.I164]*(-0.033096+0.014259*(([.F164]+273.15)/100)-0.0017*((([.F164]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G165]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J165]*112/[.J$499])*([.H165]/273)","v":0},{"$f":"of:=[.K165]/[.M165]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F165]+273.15))+143.3483*LN(([.F165]+273.15)/100)-21.8492*(([.F165]+273.15)/100)+[.I165]*(-0.033096+0.014259*(([.F165]+273.15)/100)-0.0017*((([.F165]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G168]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J168]*112/[.J$499])*([.H168]/273)","v":0},{"$f":"of:=[.K168]/[.M168]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F168]+273.15))+143.3483*LN(([.F168]+273.15)/100)-21.8492*(([.F168]+273.15)/100)+[.I168]*(-0.033096+0.014259*(([.F168]+273.15)/100)-0.0017*((([.F168]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G169]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J169]*112/[.J$499])*([.H169]/273)","v":0},{"$f":"of:=[.K169]/[.M169]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F169]+273.15))+143.3483*LN(([.F169]+273.15)/100)-21.8492*(([.F169]+273.15)/100)+[.I169]*(-0.033096+0.014259*(([.F169]+273.15)/100)-0.0017*((([.F169]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G170]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J170]*112/[.J$499])*([.H170]/273)","v":0},{"$f":"of:=[.K170]/[.M170]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F170]+273.15))+143.3483*LN(([.F170]+273.15)/100)-21.8492*(([.F170]+273.15)/100)+[.I170]*(-0.033096+0.014259*(([.F170]+273.15)/100)-0.0017*((([.F170]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G171]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J171]*112/[.J$499])*([.H171]/273)","v":0},{"$f":"of:=[.K171]/[.M171]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F171]+273.15))+143.3483*LN(([.F171]+273.15)/100)-21.8492*(([.F171]+273.15)/100)+[.I171]*(-0.033096+0.014259*(([.F171]+273.15)/100)-0.0017*((([.F171]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G172]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J172]*112/[.J$499])*([.H172]/273)","v":0},{"$f":"of:=[.K172]/[.M172]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F172]+273.15))+143.3483*LN(([.F172]+273.15)/100)-21.8492*(([.F172]+273.15)/100)+[.I172]*(-0.033096+0.014259*(([.F172]+273.15)/100)-0.0017*((([.F172]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G173]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J173]*112/[.J$499])*([.H173]/273)","v":0},{"$f":"of:=[.K173]/[.M173]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F173]+273.15))+143.3483*LN(([.F173]+273.15)/100)-21.8492*(([.F173]+273.15)/100)+[.I173]*(-0.033096+0.014259*(([.F173]+273.15)/100)-0.0017*((([.F173]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G174]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J174]*112/[.J$499])*([.H174]/273)","v":0},{"$f":"of:=[.K174]/[.M174]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F174]+273.15))+143.3483*LN(([.F174]+273.15)/100)-21.8492*(([.F174]+273.15)/100)+[.I174]*(-0.033096+0.014259*(([.F174]+273.15)/100)-0.0017*((([.F174]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G175]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J175]*112/[.J$499])*([.H175]/273)","v":0},{"$f":"of:=[.K175]/[.M175]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F175]+273.15))+143.3483*LN(([.F175]+273.15)/100)-21.8492*(([.F175]+273.15)/100)+[.I175]*(-0.033096+0.014259*(([.F175]+273.15)/100)-0.0017*((([.F175]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G176]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J176]*112/[.J$499])*([.H176]/273)","v":0},{"$f":"of:=[.K176]/[.M176]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F176]+273.15))+143.3483*LN(([.F176]+273.15)/100)-21.8492*(([.F176]+273.15)/100)+[.I176]*(-0.033096+0.014259*(([.F176]+273.15)/100)-0.0017*((([.F176]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G177]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J177]*112/[.J$499])*([.H177]/273)","v":0},{"$f":"of:=[.K177]/[.M177]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F177]+273.15))+143.3483*LN(([.F177]+273.15)/100)-21.8492*(([.F177]+273.15)/100)+[.I177]*(-0.033096+0.014259*(([.F177]+273.15)/100)-0.0017*((([.F177]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G180]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J180]*112/[.J$499])*([.H180]/273)","v":0},{"$f":"of:=[.K180]/[.M180]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F180]+273.15))+143.3483*LN(([.F180]+273.15)/100)-21.8492*(([.F180]+273.15)/100)+[.I180]*(-0.033096+0.014259*(([.F180]+273.15)/100)-0.0017*((([.F180]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G181]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J181]*112/[.J$499])*([.H181]/273)","v":0},{"$f":"of:=[.K181]/[.M181]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F181]+273.15))+143.3483*LN(([.F181]+273.15)/100)-21.8492*(([.F181]+273.15)/100)+[.I181]*(-0.033096+0.014259*(([.F181]+273.15)/100)-0.0017*((([.F181]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G182]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J182]*112/[.J$499])*([.H182]/273)","v":0},{"$f":"of:=[.K182]/[.M182]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F182]+273.15))+143.3483*LN(([.F182]+273.15)/100)-21.8492*(([.F182]+273.15)/100)+[.I182]*(-0.033096+0.014259*(([.F182]+273.15)/100)-0.0017*((([.F182]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G183]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J183]*112/[.J$499])*([.H183]/273)","v":0},{"$f":"of:=[.K183]/[.M183]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F183]+273.15))+143.3483*LN(([.F183]+273.15)/100)-21.8492*(([.F183]+273.15)/100)+[.I183]*(-0.033096+0.014259*(([.F183]+273.15)/100)-0.0017*((([.F183]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G184]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J184]*112/[.J$499])*([.H184]/273)","v":0},{"$f":"of:=[.K184]/[.M184]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F184]+273.15))+143.3483*LN(([.F184]+273.15)/100)-21.8492*(([.F184]+273.15)/100)+[.I184]*(-0.033096+0.014259*(([.F184]+273.15)/100)-0.0017*((([.F184]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G185]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J185]*112/[.J$499])*([.H185]/273)","v":0},{"$f":"of:=[.K185]/[.M185]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F185]+273.15))+143.3483*LN(([.F185]+273.15)/100)-21.8492*(([.F185]+273.15)/100)+[.I185]*(-0.033096+0.014259*(([.F185]+273.15)/100)-0.0017*((([.F185]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G186]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J186]*112/[.J$499])*([.H186]/273)","v":0},{"$f":"of:=[.K186]/[.M186]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F186]+273.15))+143.3483*LN(([.F186]+273.15)/100)-21.8492*(([.F186]+273.15)/100)+[.I186]*(-0.033096+0.014259*(([.F186]+273.15)/100)-0.0017*((([.F186]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G187]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J187]*112/[.J$499])*([.H187]/273)","v":0},{"$f":"of:=[.K187]/[.M187]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F187]+273.15))+143.3483*LN(([.F187]+273.15)/100)-21.8492*(([.F187]+273.15)/100)+[.I187]*(-0.033096+0.014259*(([.F187]+273.15)/100)-0.0017*((([.F187]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G188]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J188]*112/[.J$499])*([.H188]/273)","v":0},{"$f":"of:=[.K188]/[.M188]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F188]+273.15))+143.3483*LN(([.F188]+273.15)/100)-21.8492*(([.F188]+273.15)/100)+[.I188]*(-0.033096+0.014259*(([.F188]+273.15)/100)-0.0017*((([.F188]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G189]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J189]*112/[.J$499])*([.H189]/273)","v":0},{"$f":"of:=[.K189]/[.M189]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F189]+273.15))+143.3483*LN(([.F189]+273.15)/100)-21.8492*(([.F189]+273.15)/100)+[.I189]*(-0.033096+0.014259*(([.F189]+273.15)/100)-0.0017*((([.F189]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G192]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J192]*112/[.J$499])*([.H192]/273)","v":0},{"$f":"of:=[.K192]/[.M192]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F192]+273.15))+143.3483*LN(([.F192]+273.15)/100)-21.8492*(([.F192]+273.15)/100)+[.I192]*(-0.033096+0.014259*(([.F192]+273.15)/100)-0.0017*((([.F192]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G193]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J193]*112/[.J$499])*([.H193]/273)","v":0},{"$f":"of:=[.K193]/[.M193]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F193]+273.15))+143.3483*LN(([.F193]+273.15)/100)-21.8492*(([.F193]+273.15)/100)+[.I193]*(-0.033096+0.014259*(([.F193]+273.15)/100)-0.0017*((([.F193]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G194]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J194]*112/[.J$499])*([.H194]/273)","v":0},{"$f":"of:=[.K194]/[.M194]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F194]+273.15))+143.3483*LN(([.F194]+273.15)/100)-21.8492*(([.F194]+273.15)/100)+[.I194]*(-0.033096+0.014259*(([.F194]+273.15)/100)-0.0017*((([.F194]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G195]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J195]*112/[.J$499])*([.H195]/273)","v":0},{"$f":"of:=[.K195]/[.M195]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F195]+273.15))+143.3483*LN(([.F195]+273.15)/100)-21.8492*(([.F195]+273.15)/100)+[.I195]*(-0.033096+0.014259*(([.F195]+273.15)/100)-0.0017*((([.F195]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G196]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J196]*112/[.J$499])*([.H196]/273)","v":0},{"$f":"of:=[.K196]/[.M196]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F196]+273.15))+143.3483*LN(([.F196]+273.15)/100)-21.8492*(([.F196]+273.15)/100)+[.I196]*(-0.033096+0.014259*(([.F196]+273.15)/100)-0.0017*((([.F196]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G197]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J197]*112/[.J$499])*([.H197]/273)","v":0},{"$f":"of:=[.K197]/[.M197]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F197]+273.15))+143.3483*LN(([.F197]+273.15)/100)-21.8492*(([.F197]+273.15)/100)+[.I197]*(-0.033096+0.014259*(([.F197]+273.15)/100)-0.0017*((([.F197]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G198]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J198]*112/[.J$499])*([.H198]/273)","v":0},{"$f":"of:=[.K198]/[.M198]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F198]+273.15))+143.3483*LN(([.F198]+273.15)/100)-21.8492*(([.F198]+273.15)/100)+[.I198]*(-0.033096+0.014259*(([.F198]+273.15)/100)-0.0017*((([.F198]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G199]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J199]*112/[.J$499])*([.H199]/273)","v":0},{"$f":"of:=[.K199]/[.M199]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F199]+273.15))+143.3483*LN(([.F199]+273.15)/100)-21.8492*(([.F199]+273.15)/100)+[.I199]*(-0.033096+0.014259*(([.F199]+273.15)/100)-0.0017*((([.F199]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G200]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J200]*112/[.J$499])*([.H200]/273)","v":0},{"$f":"of:=[.K200]/[.M200]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F200]+273.15))+143.3483*LN(([.F200]+273.15)/100)-21.8492*(([.F200]+273.15)/100)+[.I200]*(-0.033096+0.014259*(([.F200]+273.15)/100)-0.0017*((([.F200]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G201]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J201]*112/[.J$499])*([.H201]/273)","v":0},{"$f":"of:=[.K201]/[.M201]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F201]+273.15))+143.3483*LN(([.F201]+273.15)/100)-21.8492*(([.F201]+273.15)/100)+[.I201]*(-0.033096+0.014259*(([.F201]+273.15)/100)-0.0017*((([.F201]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G204]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J204]*112/[.J$499])*([.H204]/273)","v":0},{"$f":"of:=[.K204]/[.M204]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F204]+273.15))+143.3483*LN(([.F204]+273.15)/100)-21.8492*(([.F204]+273.15)/100)+[.I204]*(-0.033096+0.014259*(([.F204]+273.15)/100)-0.0017*((([.F204]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G205]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J205]*112/[.J$499])*([.H205]/273)","v":0},{"$f":"of:=[.K205]/[.M205]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F205]+273.15))+143.3483*LN(([.F205]+273.15)/100)-21.8492*(([.F205]+273.15)/100)+[.I205]*(-0.033096+0.014259*(([.F205]+273.15)/100)-0.0017*((([.F205]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G206]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J206]*112/[.J$499])*([.H206]/273)","v":0},{"$f":"of:=[.K206]/[.M206]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F206]+273.15))+143.3483*LN(([.F206]+273.15)/100)-21.8492*(([.F206]+273.15)/100)+[.I206]*(-0.033096+0.014259*(([.F206]+273.15)/100)-0.0017*((([.F206]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G207]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J207]*112/[.J$499])*([.H207]/273)","v":0},{"$f":"of:=[.K207]/[.M207]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F207]+273.15))+143.3483*LN(([.F207]+273.15)/100)-21.8492*(([.F207]+273.15)/100)+[.I207]*(-0.033096+0.014259*(([.F207]+273.15)/100)-0.0017*((([.F207]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G208]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J208]*112/[.J$499])*([.H208]/273)","v":0},{"$f":"of:=[.K208]/[.M208]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F208]+273.15))+143.3483*LN(([.F208]+273.15)/100)-21.8492*(([.F208]+273.15)/100)+[.I208]*(-0.033096+0.014259*(([.F208]+273.15)/100)-0.0017*((([.F208]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G209]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J209]*112/[.J$499])*([.H209]/273)","v":0},{"$f":"of:=[.K209]/[.M209]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F209]+273.15))+143.3483*LN(([.F209]+273.15)/100)-21.8492*(([.F209]+273.15)/100)+[.I209]*(-0.033096+0.014259*(([.F209]+273.15)/100)-0.0017*((([.F209]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G210]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J210]*112/[.J$499])*([.H210]/273)","v":0},{"$f":"of:=[.K210]/[.M210]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F210]+273.15))+143.3483*LN(([.F210]+273.15)/100)-21.8492*(([.F210]+273.15)/100)+[.I210]*(-0.033096+0.014259*(([.F210]+273.15)/100)-0.0017*((([.F210]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G211]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J211]*112/[.J$499])*([.H211]/273)","v":0},{"$f":"of:=[.K211]/[.M211]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F211]+273.15))+143.3483*LN(([.F211]+273.15)/100)-21.8492*(([.F211]+273.15)/100)+[.I211]*(-0.033096+0.014259*(([.F211]+273.15)/100)-0.0017*((([.F211]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G212]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J212]*112/[.J$499])*([.H212]/273)","v":0},{"$f":"of:=[.K212]/[.M212]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F212]+273.15))+143.3483*LN(([.F212]+273.15)/100)-21.8492*(([.F212]+273.15)/100)+[.I212]*(-0.033096+0.014259*(([.F212]+273.15)/100)-0.0017*((([.F212]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G213]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J213]*112/[.J$499])*([.H213]/273)","v":0},{"$f":"of:=[.K213]/[.M213]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F213]+273.15))+143.3483*LN(([.F213]+273.15)/100)-21.8492*(([.F213]+273.15)/100)+[.I213]*(-0.033096+0.014259*(([.F213]+273.15)/100)-0.0017*((([.F213]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G216]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J216]*112/[.J$499])*([.H216]/273)","v":0},{"$f":"of:=[.K216]/[.M216]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F216]+273.15))+143.3483*LN(([.F216]+273.15)/100)-21.8492*(([.F216]+273.15)/100)+[.I216]*(-0.033096+0.014259*(([.F216]+273.15)/100)-0.0017*((([.F216]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G217]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J217]*112/[.J$499])*([.H217]/273)","v":0},{"$f":"of:=[.K217]/[.M217]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F217]+273.15))+143.3483*LN(([.F217]+273.15)/100)-21.8492*(([.F217]+273.15)/100)+[.I217]*(-0.033096+0.014259*(([.F217]+273.15)/100)-0.0017*((([.F217]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G218]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J218]*112/[.J$499])*([.H218]/273)","v":0},{"$f":"of:=[.K218]/[.M218]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F218]+273.15))+143.3483*LN(([.F218]+273.15)/100)-21.8492*(([.F218]+273.15)/100)+[.I218]*(-0.033096+0.014259*(([.F218]+273.15)/100)-0.0017*((([.F218]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G219]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J219]*112/[.J$499])*([.H219]/273)","v":0},{"$f":"of:=[.K219]/[.M219]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F219]+273.15))+143.3483*LN(([.F219]+273.15)/100)-21.8492*(([.F219]+273.15)/100)+[.I219]*(-0.033096+0.014259*(([.F219]+273.15)/100)-0.0017*((([.F219]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G220]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J220]*112/[.J$499])*([.H220]/273)","v":0},{"$f":"of:=[.K220]/[.M220]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F220]+273.15))+143.3483*LN(([.F220]+273.15)/100)-21.8492*(([.F220]+273.15)/100)+[.I220]*(-0.033096+0.014259*(([.F220]+273.15)/100)-0.0017*((([.F220]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G221]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J221]*112/[.J$499])*([.H221]/273)","v":0},{"$f":"of:=[.K221]/[.M221]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F221]+273.15))+143.3483*LN(([.F221]+273.15)/100)-21.8492*(([.F221]+273.15)/100)+[.I221]*(-0.033096+0.014259*(([.F221]+273.15)/100)-0.0017*((([.F221]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G222]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J222]*112/[.J$499])*([.H222]/273)","v":0},{"$f":"of:=[.K222]/[.M222]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F222]+273.15))+143.3483*LN(([.F222]+273.15)/100)-21.8492*(([.F222]+273.15)/100)+[.I222]*(-0.033096+0.014259*(([.F222]+273.15)/100)-0.0017*((([.F222]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G223]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J223]*112/[.J$499])*([.H223]/273)","v":0},{"$f":"of:=[.K223]/[.M223]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F223]+273.15))+143.3483*LN(([.F223]+273.15)/100)-21.8492*(([.F223]+273.15)/100)+[.I223]*(-0.033096+0.014259*(([.F223]+273.15)/100)-0.0017*((([.F223]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G224]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J224]*112/[.J$499])*([.H224]/273)","v":0},{"$f":"of:=[.K224]/[.M224]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F224]+273.15))+143.3483*LN(([.F224]+273.15)/100)-21.8492*(([.F224]+273.15)/100)+[.I224]*(-0.033096+0.014259*(([.F224]+273.15)/100)-0.0017*((([.F224]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G225]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J225]*112/[.J$499])*([.H225]/273)","v":0},{"$f":"of:=[.K225]/[.M225]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F225]+273.15))+143.3483*LN(([.F225]+273.15)/100)-21.8492*(([.F225]+273.15)/100)+[.I225]*(-0.033096+0.014259*(([.F225]+273.15)/100)-0.0017*((([.F225]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G228]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J228]*112/[.J$499])*([.H228]/273)","v":0},{"$f":"of:=[.K228]/[.M228]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F228]+273.15))+143.3483*LN(([.F228]+273.15)/100)-21.8492*(([.F228]+273.15)/100)+[.I228]*(-0.033096+0.014259*(([.F228]+273.15)/100)-0.0017*((([.F228]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G229]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J229]*112/[.J$499])*([.H229]/273)","v":0},{"$f":"of:=[.K229]/[.M229]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F229]+273.15))+143.3483*LN(([.F229]+273.15)/100)-21.8492*(([.F229]+273.15)/100)+[.I229]*(-0.033096+0.014259*(([.F229]+273.15)/100)-0.0017*((([.F229]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G230]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J230]*112/[.J$499])*([.H230]/273)","v":0},{"$f":"of:=[.K230]/[.M230]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F230]+273.15))+143.3483*LN(([.F230]+273.15)/100)-21.8492*(([.F230]+273.15)/100)+[.I230]*(-0.033096+0.014259*(([.F230]+273.15)/100)-0.0017*((([.F230]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G231]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J231]*112/[.J$499])*([.H231]/273)","v":0},{"$f":"of:=[.K231]/[.M231]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F231]+273.15))+143.3483*LN(([.F231]+273.15)/100)-21.8492*(([.F231]+273.15)/100)+[.I231]*(-0.033096+0.014259*(([.F231]+273.15)/100)-0.0017*((([.F231]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G232]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J232]*112/[.J$499])*([.H232]/273)","v":0},{"$f":"of:=[.K232]/[.M232]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F232]+273.15))+143.3483*LN(([.F232]+273.15)/100)-21.8492*(([.F232]+273.15)/100)+[.I232]*(-0.033096+0.014259*(([.F232]+273.15)/100)-0.0017*((([.F232]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G233]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J233]*112/[.J$499])*([.H233]/273)","v":0},{"$f":"of:=[.K233]/[.M233]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F233]+273.15))+143.3483*LN(([.F233]+273.15)/100)-21.8492*(([.F233]+273.15)/100)+[.I233]*(-0.033096+0.014259*(([.F233]+273.15)/100)-0.0017*((([.F233]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G234]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J234]*112/[.J$499])*([.H234]/273)","v":0},{"$f":"of:=[.K234]/[.M234]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F234]+273.15))+143.3483*LN(([.F234]+273.15)/100)-21.8492*(([.F234]+273.15)/100)+[.I234]*(-0.033096+0.014259*(([.F234]+273.15)/100)-0.0017*((([.F234]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G235]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J235]*112/[.J$499])*([.H235]/273)","v":0},{"$f":"of:=[.K235]/[.M235]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F235]+273.15))+143.3483*LN(([.F235]+273.15)/100)-21.8492*(([.F235]+273.15)/100)+[.I235]*(-0.033096+0.014259*(([.F235]+273.15)/100)-0.0017*((([.F235]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G236]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J236]*112/[.J$499])*([.H236]/273)","v":0},{"$f":"of:=[.K236]/[.M236]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F236]+273.15))+143.3483*LN(([.F236]+273.15)/100)-21.8492*(([.F236]+273.15)/100)+[.I236]*(-0.033096+0.014259*(([.F236]+273.15)/100)-0.0017*((([.F236]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G237]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J237]*112/[.J$499])*([.H237]/273)","v":0},{"$f":"of:=[.K237]/[.M237]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F237]+273.15))+143.3483*LN(([.F237]+273.15)/100)-21.8492*(([.F237]+273.15)/100)+[.I237]*(-0.033096+0.014259*(([.F237]+273.15)/100)-0.0017*((([.F237]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G240]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J240]*112/[.J$499])*([.H240]/273)","v":0},{"$f":"of:=[.K240]/[.M240]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F240]+273.15))+143.3483*LN(([.F240]+273.15)/100)-21.8492*(([.F240]+273.15)/100)+[.I240]*(-0.033096+0.014259*(([.F240]+273.15)/100)-0.0017*((([.F240]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G241]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J241]*112/[.J$499])*([.H241]/273)","v":0},{"$f":"of:=[.K241]/[.M241]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F241]+273.15))+143.3483*LN(([.F241]+273.15)/100)-21.8492*(([.F241]+273.15)/100)+[.I241]*(-0.033096+0.014259*(([.F241]+273.15)/100)-0.0017*((([.F241]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G242]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J242]*112/[.J$499])*([.H242]/273)","v":0},{"$f":"of:=[.K242]/[.M242]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F242]+273.15))+143.3483*LN(([.F242]+273.15)/100)-21.8492*(([.F242]+273.15)/100)+[.I242]*(-0.033096+0.014259*(([.F242]+273.15)/100)-0.0017*((([.F242]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G243]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J243]*112/[.J$499])*([.H243]/273)","v":0},{"$f":"of:=[.K243]/[.M243]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F243]+273.15))+143.3483*LN(([.F243]+273.15)/100)-21.8492*(([.F243]+273.15)/100)+[.I243]*(-0.033096+0.014259*(([.F243]+273.15)/100)-0.0017*((([.F243]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G244]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J244]*112/[.J$499])*([.H244]/273)","v":0},{"$f":"of:=[.K244]/[.M244]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F244]+273.15))+143.3483*LN(([.F244]+273.15)/100)-21.8492*(([.F244]+273.15)/100)+[.I244]*(-0.033096+0.014259*(([.F244]+273.15)/100)-0.0017*((([.F244]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G245]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J245]*112/[.J$499])*([.H245]/273)","v":0},{"$f":"of:=[.K245]/[.M245]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F245]+273.15))+143.3483*LN(([.F245]+273.15)/100)-21.8492*(([.F245]+273.15)/100)+[.I245]*(-0.033096+0.014259*(([.F245]+273.15)/100)-0.0017*((([.F245]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G246]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J246]*112/[.J$499])*([.H246]/273)","v":0},{"$f":"of:=[.K246]/[.M246]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F246]+273.15))+143.3483*LN(([.F246]+273.15)/100)-21.8492*(([.F246]+273.15)/100)+[.I246]*(-0.033096+0.014259*(([.F246]+273.15)/100)-0.0017*((([.F246]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G247]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J247]*112/[.J$499])*([.H247]/273)","v":0},{"$f":"of:=[.K247]/[.M247]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F247]+273.15))+143.3483*LN(([.F247]+273.15)/100)-21.8492*(([.F247]+273.15)/100)+[.I247]*(-0.033096+0.014259*(([.F247]+273.15)/100)-0.0017*((([.F247]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G248]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J248]*112/[.J$499])*([.H248]/273)","v":0},{"$f":"of:=[.K248]/[.M248]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F248]+273.15))+143.3483*LN(([.F248]+273.15)/100)-21.8492*(([.F248]+273.15)/100)+[.I248]*(-0.033096+0.014259*(([.F248]+273.15)/100)-0.0017*((([.F248]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G249]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J249]*112/[.J$499])*([.H249]/273)","v":0},{"$f":"of:=[.K249]/[.M249]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F249]+273.15))+143.3483*LN(([.F249]+273.15)/100)-21.8492*(([.F249]+273.15)/100)+[.I249]*(-0.033096+0.014259*(([.F249]+273.15)/100)-0.0017*((([.F249]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G252]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J252]*112/[.J$499])*([.H252]/273)","v":0},{"$f":"of:=[.K252]/[.M252]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F252]+273.15))+143.3483*LN(([.F252]+273.15)/100)-21.8492*(([.F252]+273.15)/100)+[.I252]*(-0.033096+0.014259*(([.F252]+273.15)/100)-0.0017*((([.F252]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G253]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J253]*112/[.J$499])*([.H253]/273)","v":0},{"$f":"of:=[.K253]/[.M253]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F253]+273.15))+143.3483*LN(([.F253]+273.15)/100)-21.8492*(([.F253]+273.15)/100)+[.I253]*(-0.033096+0.014259*(([.F253]+273.15)/100)-0.0017*((([.F253]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G254]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J254]*112/[.J$499])*([.H254]/273)","v":0},{"$f":"of:=[.K254]/[.M254]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F254]+273.15))+143.3483*LN(([.F254]+273.15)/100)-21.8492*(([.F254]+273.15)/100)+[.I254]*(-0.033096+0.014259*(([.F254]+273.15)/100)-0.0017*((([.F254]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G255]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J255]*112/[.J$499])*([.H255]/273)","v":0},{"$f":"of:=[.K255]/[.M255]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F255]+273.15))+143.3483*LN(([.F255]+273.15)/100)-21.8492*(([.F255]+273.15)/100)+[.I255]*(-0.033096+0.014259*(([.F255]+273.15)/100)-0.0017*((([.F255]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G256]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J256]*112/[.J$499])*([.H256]/273)","v":0},{"$f":"of:=[.K256]/[.M256]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F256]+273.15))+143.3483*LN(([.F256]+273.15)/100)-21.8492*(([.F256]+273.15)/100)+[.I256]*(-0.033096+0.014259*(([.F256]+273.15)/100)-0.0017*((([.F256]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G257]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J257]*112/[.J$499])*([.H257]/273)","v":0},{"$f":"of:=[.K257]/[.M257]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F257]+273.15))+143.3483*LN(([.F257]+273.15)/100)-21.8492*(([.F257]+273.15)/100)+[.I257]*(-0.033096+0.014259*(([.F257]+273.15)/100)-0.0017*((([.F257]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G258]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J258]*112/[.J$499])*([.H258]/273)","v":0},{"$f":"of:=[.K258]/[.M258]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F258]+273.15))+143.3483*LN(([.F258]+273.15)/100)-21.8492*(([.F258]+273.15)/100)+[.I258]*(-0.033096+0.014259*(([.F258]+273.15)/100)-0.0017*((([.F258]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G259]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J259]*112/[.J$499])*([.H259]/273)","v":0},{"$f":"of:=[.K259]/[.M259]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F259]+273.15))+143.3483*LN(([.F259]+273.15)/100)-21.8492*(([.F259]+273.15)/100)+[.I259]*(-0.033096+0.014259*(([.F259]+273.15)/100)-0.0017*((([.F259]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G260]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J260]*112/[.J$499])*([.H260]/273)","v":0},{"$f":"of:=[.K260]/[.M260]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F260]+273.15))+143.3483*LN(([.F260]+273.15)/100)-21.8492*(([.F260]+273.15)/100)+[.I260]*(-0.033096+0.014259*(([.F260]+273.15)/100)-0.0017*((([.F260]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G261]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J261]*112/[.J$499])*([.H261]/273)","v":0},{"$f":"of:=[.K261]/[.M261]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F261]+273.15))+143.3483*LN(([.F261]+273.15)/100)-21.8492*(([.F261]+273.15)/100)+[.I261]*(-0.033096+0.014259*(([.F261]+273.15)/100)-0.0017*((([.F261]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G264]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J264]*112/[.J$499])*([.H264]/273)","v":0},{"$f":"of:=[.K264]/[.M264]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F264]+273.15))+143.3483*LN(([.F264]+273.15)/100)-21.8492*(([.F264]+273.15)/100)+[.I264]*(-0.033096+0.014259*(([.F264]+273.15)/100)-0.0017*((([.F264]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G265]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J265]*112/[.J$499])*([.H265]/273)","v":0},{"$f":"of:=[.K265]/[.M265]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F265]+273.15))+143.3483*LN(([.F265]+273.15)/100)-21.8492*(([.F265]+273.15)/100)+[.I265]*(-0.033096+0.014259*(([.F265]+273.15)/100)-0.0017*((([.F265]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G266]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J266]*112/[.J$499])*([.H266]/273)","v":0},{"$f":"of:=[.K266]/[.M266]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F266]+273.15))+143.3483*LN(([.F266]+273.15)/100)-21.8492*(([.F266]+273.15)/100)+[.I266]*(-0.033096+0.014259*(([.F266]+273.15)/100)-0.0017*((([.F266]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G267]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J267]*112/[.J$499])*([.H267]/273)","v":0},{"$f":"of:=[.K267]/[.M267]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F267]+273.15))+143.3483*LN(([.F267]+273.15)/100)-21.8492*(([.F267]+273.15)/100)+[.I267]*(-0.033096+0.014259*(([.F267]+273.15)/100)-0.0017*((([.F267]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G268]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J268]*112/[.J$499])*([.H268]/273)","v":0},{"$f":"of:=[.K268]/[.M268]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F268]+273.15))+143.3483*LN(([.F268]+273.15)/100)-21.8492*(([.F268]+273.15)/100)+[.I268]*(-0.033096+0.014259*(([.F268]+273.15)/100)-0.0017*((([.F268]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G269]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J269]*112/[.J$499])*([.H269]/273)","v":0},{"$f":"of:=[.K269]/[.M269]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F269]+273.15))+143.3483*LN(([.F269]+273.15)/100)-21.8492*(([.F269]+273.15)/100)+[.I269]*(-0.033096+0.014259*(([.F269]+273.15)/100)-0.0017*((([.F269]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G270]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J270]*112/[.J$499])*([.H270]/273)","v":0},{"$f":"of:=[.K270]/[.M270]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F270]+273.15))+143.3483*LN(([.F270]+273.15)/100)-21.8492*(([.F270]+273.15)/100)+[.I270]*(-0.033096+0.014259*(([.F270]+273.15)/100)-0.0017*((([.F270]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G271]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J271]*112/[.J$499])*([.H271]/273)","v":0},{"$f":"of:=[.K271]/[.M271]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F271]+273.15))+143.3483*LN(([.F271]+273.15)/100)-21.8492*(([.F271]+273.15)/100)+[.I271]*(-0.033096+0.014259*(([.F271]+273.15)/100)-0.0017*((([.F271]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G272]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J272]*112/[.J$499])*([.H272]/273)","v":0},{"$f":"of:=[.K272]/[.M272]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F272]+273.15))+143.3483*LN(([.F272]+273.15)/100)-21.8492*(([.F272]+273.15)/100)+[.I272]*(-0.033096+0.014259*(([.F272]+273.15)/100)-0.0017*((([.F272]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G273]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J273]*112/[.J$499])*([.H273]/273)","v":0},{"$f":"of:=[.K273]/[.M273]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F273]+273.15))+143.3483*LN(([.F273]+273.15)/100)-21.8492*(([.F273]+273.15)/100)+[.I273]*(-0.033096+0.014259*(([.F273]+273.15)/100)-0.0017*((([.F273]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G276]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J276]*112/[.J$499])*([.H276]/273)","v":0},{"$f":"of:=[.K276]/[.M276]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F276]+273.15))+143.3483*LN(([.F276]+273.15)/100)-21.8492*(([.F276]+273.15)/100)+[.I276]*(-0.033096+0.014259*(([.F276]+273.15)/100)-0.0017*((([.F276]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G277]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J277]*112/[.J$499])*([.H277]/273)","v":0},{"$f":"of:=[.K277]/[.M277]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F277]+273.15))+143.3483*LN(([.F277]+273.15)/100)-21.8492*(([.F277]+273.15)/100)+[.I277]*(-0.033096+0.014259*(([.F277]+273.15)/100)-0.0017*((([.F277]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G278]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J278]*112/[.J$499])*([.H278]/273)","v":0},{"$f":"of:=[.K278]/[.M278]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F278]+273.15))+143.3483*LN(([.F278]+273.15)/100)-21.8492*(([.F278]+273.15)/100)+[.I278]*(-0.033096+0.014259*(([.F278]+273.15)/100)-0.0017*((([.F278]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G279]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J279]*112/[.J$499])*([.H279]/273)","v":0},{"$f":"of:=[.K279]/[.M279]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F279]+273.15))+143.3483*LN(([.F279]+273.15)/100)-21.8492*(([.F279]+273.15)/100)+[.I279]*(-0.033096+0.014259*(([.F279]+273.15)/100)-0.0017*((([.F279]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G280]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J280]*112/[.J$499])*([.H280]/273)","v":0},{"$f":"of:=[.K280]/[.M280]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F280]+273.15))+143.3483*LN(([.F280]+273.15)/100)-21.8492*(([.F280]+273.15)/100)+[.I280]*(-0.033096+0.014259*(([.F280]+273.15)/100)-0.0017*((([.F280]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G281]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J281]*112/[.J$499])*([.H281]/273)","v":0},{"$f":"of:=[.K281]/[.M281]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F281]+273.15))+143.3483*LN(([.F281]+273.15)/100)-21.8492*(([.F281]+273.15)/100)+[.I281]*(-0.033096+0.014259*(([.F281]+273.15)/100)-0.0017*((([.F281]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G282]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J282]*112/[.J$499])*([.H282]/273)","v":0},{"$f":"of:=[.K282]/[.M282]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F282]+273.15))+143.3483*LN(([.F282]+273.15)/100)-21.8492*(([.F282]+273.15)/100)+[.I282]*(-0.033096+0.014259*(([.F282]+273.15)/100)-0.0017*((([.F282]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G283]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J283]*112/[.J$499])*([.H283]/273)","v":0},{"$f":"of:=[.K283]/[.M283]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F283]+273.15))+143.3483*LN(([.F283]+273.15)/100)-21.8492*(([.F283]+273.15)/100)+[.I283]*(-0.033096+0.014259*(([.F283]+273.15)/100)-0.0017*((([.F283]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G284]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J284]*112/[.J$499])*([.H284]/273)","v":0},{"$f":"of:=[.K284]/[.M284]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F284]+273.15))+143.3483*LN(([.F284]+273.15)/100)-21.8492*(([.F284]+273.15)/100)+[.I284]*(-0.033096+0.014259*(([.F284]+273.15)/100)-0.0017*((([.F284]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G285]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J285]*112/[.J$499])*([.H285]/273)","v":0},{"$f":"of:=[.K285]/[.M285]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F285]+273.15))+143.3483*LN(([.F285]+273.15)/100)-21.8492*(([.F285]+273.15)/100)+[.I285]*(-0.033096+0.014259*(([.F285]+273.15)/100)-0.0017*((([.F285]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G288]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J288]*112/[.J$499])*([.H288]/273)","v":0},{"$f":"of:=[.K288]/[.M288]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F288]+273.15))+143.3483*LN(([.F288]+273.15)/100)-21.8492*(([.F288]+273.15)/100)+[.I288]*(-0.033096+0.014259*(([.F288]+273.15)/100)-0.0017*((([.F288]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G289]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J289]*112/[.J$499])*([.H289]/273)","v":0},{"$f":"of:=[.K289]/[.M289]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F289]+273.15))+143.3483*LN(([.F289]+273.15)/100)-21.8492*(([.F289]+273.15)/100)+[.I289]*(-0.033096+0.014259*(([.F289]+273.15)/100)-0.0017*((([.F289]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G290]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J290]*112/[.J$499])*([.H290]/273)","v":0},{"$f":"of:=[.K290]/[.M290]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F290]+273.15))+143.3483*LN(([.F290]+273.15)/100)-21.8492*(([.F290]+273.15)/100)+[.I290]*(-0.033096+0.014259*(([.F290]+273.15)/100)-0.0017*((([.F290]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G291]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J291]*112/[.J$499])*([.H291]/273)","v":0},{"$f":"of:=[.K291]/[.M291]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F291]+273.15))+143.3483*LN(([.F291]+273.15)/100)-21.8492*(([.F291]+273.15)/100)+[.I291]*(-0.033096+0.014259*(([.F291]+273.15)/100)-0.0017*((([.F291]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G292]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J292]*112/[.J$499])*([.H292]/273)","v":0},{"$f":"of:=[.K292]/[.M292]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F292]+273.15))+143.3483*LN(([.F292]+273.15)/100)-21.8492*(([.F292]+273.15)/100)+[.I292]*(-0.033096+0.014259*(([.F292]+273.15)/100)-0.0017*((([.F292]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G293]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J293]*112/[.J$499])*([.H293]/273)","v":0},{"$f":"of:=[.K293]/[.M293]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F293]+273.15))+143.3483*LN(([.F293]+273.15)/100)-21.8492*(([.F293]+273.15)/100)+[.I293]*(-0.033096+0.014259*(([.F293]+273.15)/100)-0.0017*((([.F293]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G294]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J294]*112/[.J$499])*([.H294]/273)","v":0},{"$f":"of:=[.K294]/[.M294]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F294]+273.15))+143.3483*LN(([.F294]+273.15)/100)-21.8492*(([.F294]+273.15)/100)+[.I294]*(-0.033096+0.014259*(([.F294]+273.15)/100)-0.0017*((([.F294]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G295]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J295]*112/[.J$499])*([.H295]/273)","v":0},{"$f":"of:=[.K295]/[.M295]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F295]+273.15))+143.3483*LN(([.F295]+273.15)/100)-21.8492*(([.F295]+273.15)/100)+[.I295]*(-0.033096+0.014259*(([.F295]+273.15)/100)-0.0017*((([.F295]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G296]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J296]*112/[.J$499])*([.H296]/273)","v":0},{"$f":"of:=[.K296]/[.M296]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F296]+273.15))+143.3483*LN(([.F296]+273.15)/100)-21.8492*(([.F296]+273.15)/100)+[.I296]*(-0.033096+0.014259*(([.F296]+273.15)/100)-0.0017*((([.F296]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G297]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J297]*112/[.J$499])*([.H297]/273)","v":0},{"$f":"of:=[.K297]/[.M297]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F297]+273.15))+143.3483*LN(([.F297]+273.15)/100)-21.8492*(([.F297]+273.15)/100)+[.I297]*(-0.033096+0.014259*(([.F297]+273.15)/100)-0.0017*((([.F297]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G300]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J300]*112/[.J$499])*([.H300]/273)","v":0},{"$f":"of:=[.K300]/[.M300]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F300]+273.15))+143.3483*LN(([.F300]+273.15)/100)-21.8492*(([.F300]+273.15)/100)+[.I300]*(-0.033096+0.014259*(([.F300]+273.15)/100)-0.0017*((([.F300]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G301]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J301]*112/[.J$499])*([.H301]/273)","v":0},{"$f":"of:=[.K301]/[.M301]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F301]+273.15))+143.3483*LN(([.F301]+273.15)/100)-21.8492*(([.F301]+273.15)/100)+[.I301]*(-0.033096+0.014259*(([.F301]+273.15)/100)-0.0017*((([.F301]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G302]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J302]*112/[.J$499])*([.H302]/273)","v":0},{"$f":"of:=[.K302]/[.M302]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F302]+273.15))+143.3483*LN(([.F302]+273.15)/100)-21.8492*(([.F302]+273.15)/100)+[.I302]*(-0.033096+0.014259*(([.F302]+273.15)/100)-0.0017*((([.F302]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G303]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J303]*112/[.J$499])*([.H303]/273)","v":0},{"$f":"of:=[.K303]/[.M303]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F303]+273.15))+143.3483*LN(([.F303]+273.15)/100)-21.8492*(([.F303]+273.15)/100)+[.I303]*(-0.033096+0.014259*(([.F303]+273.15)/100)-0.0017*((([.F303]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G304]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J304]*112/[.J$499])*([.H304]/273)","v":0},{"$f":"of:=[.K304]/[.M304]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F304]+273.15))+143.3483*LN(([.F304]+273.15)/100)-21.8492*(([.F304]+273.15)/100)+[.I304]*(-0.033096+0.014259*(([.F304]+273.15)/100)-0.0017*((([.F304]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G305]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J305]*112/[.J$499])*([.H305]/273)","v":0},{"$f":"of:=[.K305]/[.M305]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F305]+273.15))+143.3483*LN(([.F305]+273.15)/100)-21.8492*(([.F305]+273.15)/100)+[.I305]*(-0.033096+0.014259*(([.F305]+273.15)/100)-0.0017*((([.F305]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G306]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J306]*112/[.J$499])*([.H306]/273)","v":0},{"$f":"of:=[.K306]/[.M306]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F306]+273.15))+143.3483*LN(([.F306]+273.15)/100)-21.8492*(([.F306]+273.15)/100)+[.I306]*(-0.033096+0.014259*(([.F306]+273.15)/100)-0.0017*((([.F306]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G307]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J307]*112/[.J$499])*([.H307]/273)","v":0},{"$f":"of:=[.K307]/[.M307]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F307]+273.15))+143.3483*LN(([.F307]+273.15)/100)-21.8492*(([.F307]+273.15)/100)+[.I307]*(-0.033096+0.014259*(([.F307]+273.15)/100)-0.0017*((([.F307]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G308]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J308]*112/[.J$499])*([.H308]/273)","v":0},{"$f":"of:=[.K308]/[.M308]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F308]+273.15))+143.3483*LN(([.F308]+273.15)/100)-21.8492*(([.F308]+273.15)/100)+[.I308]*(-0.033096+0.014259*(([.F308]+273.15)/100)-0.0017*((([.F308]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G309]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J309]*112/[.J$499])*([.H309]/273)","v":0},{"$f":"of:=[.K309]/[.M309]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F309]+273.15))+143.3483*LN(([.F309]+273.15)/100)-21.8492*(([.F309]+273.15)/100)+[.I309]*(-0.033096+0.014259*(([.F309]+273.15)/100)-0.0017*((([.F309]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G312]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J312]*112/[.J$499])*([.H312]/273)","v":0},{"$f":"of:=[.K312]/[.M312]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F312]+273.15))+143.3483*LN(([.F312]+273.15)/100)-21.8492*(([.F312]+273.15)/100)+[.I312]*(-0.033096+0.014259*(([.F312]+273.15)/100)-0.0017*((([.F312]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G313]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J313]*112/[.J$499])*([.H313]/273)","v":0},{"$f":"of:=[.K313]/[.M313]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F313]+273.15))+143.3483*LN(([.F313]+273.15)/100)-21.8492*(([.F313]+273.15)/100)+[.I313]*(-0.033096+0.014259*(([.F313]+273.15)/100)-0.0017*((([.F313]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G314]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J314]*112/[.J$499])*([.H314]/273)","v":0},{"$f":"of:=[.K314]/[.M314]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F314]+273.15))+143.3483*LN(([.F314]+273.15)/100)-21.8492*(([.F314]+273.15)/100)+[.I314]*(-0.033096+0.014259*(([.F314]+273.15)/100)-0.0017*((([.F314]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G315]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J315]*112/[.J$499])*([.H315]/273)","v":0},{"$f":"of:=[.K315]/[.M315]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F315]+273.15))+143.3483*LN(([.F315]+273.15)/100)-21.8492*(([.F315]+273.15)/100)+[.I315]*(-0.033096+0.014259*(([.F315]+273.15)/100)-0.0017*((([.F315]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G316]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J316]*112/[.J$499])*([.H316]/273)","v":0},{"$f":"of:=[.K316]/[.M316]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F316]+273.15))+143.3483*LN(([.F316]+273.15)/100)-21.8492*(([.F316]+273.15)/100)+[.I316]*(-0.033096+0.014259*(([.F316]+273.15)/100)-0.0017*((([.F316]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G317]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J317]*112/[.J$499])*([.H317]/273)","v":0},{"$f":"of:=[.K317]/[.M317]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F317]+273.15))+143.3483*LN(([.F317]+273.15)/100)-21.8492*(([.F317]+273.15)/100)+[.I317]*(-0.033096+0.014259*(([.F317]+273.15)/100)-0.0017*((([.F317]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G318]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J318]*112/[.J$499])*([.H318]/273)","v":0},{"$f":"of:=[.K318]/[.M318]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F318]+273.15))+143.3483*LN(([.F318]+273.15)/100)-21.8492*(([.F318]+273.15)/100)+[.I318]*(-0.033096+0.014259*(([.F318]+273.15)/100)-0.0017*((([.F318]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G319]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J319]*112/[.J$499])*([.H319]/273)","v":0},{"$f":"of:=[.K319]/[.M319]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F319]+273.15))+143.3483*LN(([.F319]+273.15)/100)-21.8492*(([.F319]+273.15)/100)+[.I319]*(-0.033096+0.014259*(([.F319]+273.15)/100)-0.0017*((([.F319]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G320]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J320]*112/[.J$499])*([.H320]/273)","v":0},{"$f":"of:=[.K320]/[.M320]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F320]+273.15))+143.3483*LN(([.F320]+273.15)/100)-21.8492*(([.F320]+273.15)/100)+[.I320]*(-0.033096+0.014259*(([.F320]+273.15)/100)-0.0017*((([.F320]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G321]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J321]*112/[.J$499])*([.H321]/273)","v":0},{"$f":"of:=[.K321]/[.M321]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F321]+273.15))+143.3483*LN(([.F321]+273.15)/100)-21.8492*(([.F321]+273.15)/100)+[.I321]*(-0.033096+0.014259*(([.F321]+273.15)/100)-0.0017*((([.F321]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G324]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J324]*112/[.J$499])*([.H324]/273)","v":0},{"$f":"of:=[.K324]/[.M324]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F324]+273.15))+143.3483*LN(([.F324]+273.15)/100)-21.8492*(([.F324]+273.15)/100)+[.I324]*(-0.033096+0.014259*(([.F324]+273.15)/100)-0.0017*((([.F324]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G325]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J325]*112/[.J$499])*([.H325]/273)","v":0},{"$f":"of:=[.K325]/[.M325]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F325]+273.15))+143.3483*LN(([.F325]+273.15)/100)-21.8492*(([.F325]+273.15)/100)+[.I325]*(-0.033096+0.014259*(([.F325]+273.15)/100)-0.0017*((([.F325]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G326]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J326]*112/[.J$499])*([.H326]/273)","v":0},{"$f":"of:=[.K326]/[.M326]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F326]+273.15))+143.3483*LN(([.F326]+273.15)/100)-21.8492*(([.F326]+273.15)/100)+[.I326]*(-0.033096+0.014259*(([.F326]+273.15)/100)-0.0017*((([.F326]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G327]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J327]*112/[.J$499])*([.H327]/273)","v":0},{"$f":"of:=[.K327]/[.M327]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F327]+273.15))+143.3483*LN(([.F327]+273.15)/100)-21.8492*(([.F327]+273.15)/100)+[.I327]*(-0.033096+0.014259*(([.F327]+273.15)/100)-0.0017*((([.F327]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G328]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J328]*112/[.J$499])*([.H328]/273)","v":0},{"$f":"of:=[.K328]/[.M328]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F328]+273.15))+143.3483*LN(([.F328]+273.15)/100)-21.8492*(([.F328]+273.15)/100)+[.I328]*(-0.033096+0.014259*(([.F328]+273.15)/100)-0.0017*((([.F328]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G329]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J329]*112/[.J$499])*([.H329]/273)","v":0},{"$f":"of:=[.K329]/[.M329]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F329]+273.15))+143.3483*LN(([.F329]+273.15)/100)-21.8492*(([.F329]+273.15)/100)+[.I329]*(-0.033096+0.014259*(([.F329]+273.15)/100)-0.0017*((([.F329]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G330]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J330]*112/[.J$499])*([.H330]/273)","v":0},{"$f":"of:=[.K330]/[.M330]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F330]+273.15))+143.3483*LN(([.F330]+273.15)/100)-21.8492*(([.F330]+273.15)/100)+[.I330]*(-0.033096+0.014259*(([.F330]+273.15)/100)-0.0017*((([.F330]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G331]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J331]*112/[.J$499])*([.H331]/273)","v":0},{"$f":"of:=[.K331]/[.M331]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F331]+273.15))+143.3483*LN(([.F331]+273.15)/100)-21.8492*(([.F331]+273.15)/100)+[.I331]*(-0.033096+0.014259*(([.F331]+273.15)/100)-0.0017*((([.F331]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G332]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J332]*112/[.J$499])*([.H332]/273)","v":0},{"$f":"of:=[.K332]/[.M332]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F332]+273.15))+143.3483*LN(([.F332]+273.15)/100)-21.8492*(([.F332]+273.15)/100)+[.I332]*(-0.033096+0.014259*(([.F332]+273.15)/100)-0.0017*((([.F332]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G333]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J333]*112/[.J$499])*([.H333]/273)","v":0},{"$f":"of:=[.K333]/[.M333]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F333]+273.15))+143.3483*LN(([.F333]+273.15)/100)-21.8492*(([.F333]+273.15)/100)+[.I333]*(-0.033096+0.014259*(([.F333]+273.15)/100)-0.0017*((([.F333]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G336]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J336]*112/[.J$499])*([.H336]/273)","v":0},{"$f":"of:=[.K336]/[.M336]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F336]+273.15))+143.3483*LN(([.F336]+273.15)/100)-21.8492*(([.F336]+273.15)/100)+[.I336]*(-0.033096+0.014259*(([.F336]+273.15)/100)-0.0017*((([.F336]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G337]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J337]*112/[.J$499])*([.H337]/273)","v":0},{"$f":"of:=[.K337]/[.M337]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F337]+273.15))+143.3483*LN(([.F337]+273.15)/100)-21.8492*(([.F337]+273.15)/100)+[.I337]*(-0.033096+0.014259*(([.F337]+273.15)/100)-0.0017*((([.F337]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G338]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J338]*112/[.J$499])*([.H338]/273)","v":0},{"$f":"of:=[.K338]/[.M338]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F338]+273.15))+143.3483*LN(([.F338]+273.15)/100)-21.8492*(([.F338]+273.15)/100)+[.I338]*(-0.033096+0.014259*(([.F338]+273.15)/100)-0.0017*((([.F338]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G339]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J339]*112/[.J$499])*([.H339]/273)","v":0},{"$f":"of:=[.K339]/[.M339]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F339]+273.15))+143.3483*LN(([.F339]+273.15)/100)-21.8492*(([.F339]+273.15)/100)+[.I339]*(-0.033096+0.014259*(([.F339]+273.15)/100)-0.0017*((([.F339]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G340]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J340]*112/[.J$499])*([.H340]/273)","v":0},{"$f":"of:=[.K340]/[.M340]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F340]+273.15))+143.3483*LN(([.F340]+273.15)/100)-21.8492*(([.F340]+273.15)/100)+[.I340]*(-0.033096+0.014259*(([.F340]+273.15)/100)-0.0017*((([.F340]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G341]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J341]*112/[.J$499])*([.H341]/273)","v":0},{"$f":"of:=[.K341]/[.M341]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F341]+273.15))+143.3483*LN(([.F341]+273.15)/100)-21.8492*(([.F341]+273.15)/100)+[.I341]*(-0.033096+0.014259*(([.F341]+273.15)/100)-0.0017*((([.F341]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G342]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J342]*112/[.J$499])*([.H342]/273)","v":0},{"$f":"of:=[.K342]/[.M342]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F342]+273.15))+143.3483*LN(([.F342]+273.15)/100)-21.8492*(([.F342]+273.15)/100)+[.I342]*(-0.033096+0.014259*(([.F342]+273.15)/100)-0.0017*((([.F342]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G343]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J343]*112/[.J$499])*([.H343]/273)","v":0},{"$f":"of:=[.K343]/[.M343]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F343]+273.15))+143.3483*LN(([.F343]+273.15)/100)-21.8492*(([.F343]+273.15)/100)+[.I343]*(-0.033096+0.014259*(([.F343]+273.15)/100)-0.0017*((([.F343]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G344]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J344]*112/[.J$499])*([.H344]/273)","v":0},{"$f":"of:=[.K344]/[.M344]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F344]+273.15))+143.3483*LN(([.F344]+273.15)/100)-21.8492*(([.F344]+273.15)/100)+[.I344]*(-0.033096+0.014259*(([.F344]+273.15)/100)-0.0017*((([.F344]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G345]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J345]*112/[.J$499])*([.H345]/273)","v":0},{"$f":"of:=[.K345]/[.M345]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F345]+273.15))+143.3483*LN(([.F345]+273.15)/100)-21.8492*(([.F345]+273.15)/100)+[.I345]*(-0.033096+0.014259*(([.F345]+273.15)/100)-0.0017*((([.F345]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G348]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J348]*112/[.J$499])*([.H348]/273)","v":0},{"$f":"of:=[.K348]/[.M348]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F348]+273.15))+143.3483*LN(([.F348]+273.15)/100)-21.8492*(([.F348]+273.15)/100)+[.I348]*(-0.033096+0.014259*(([.F348]+273.15)/100)-0.0017*((([.F348]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G349]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J349]*112/[.J$499])*([.H349]/273)","v":0},{"$f":"of:=[.K349]/[.M349]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F349]+273.15))+143.3483*LN(([.F349]+273.15)/100)-21.8492*(([.F349]+273.15)/100)+[.I349]*(-0.033096+0.014259*(([.F349]+273.15)/100)-0.0017*((([.F349]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G350]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J350]*112/[.J$499])*([.H350]/273)","v":0},{"$f":"of:=[.K350]/[.M350]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F350]+273.15))+143.3483*LN(([.F350]+273.15)/100)-21.8492*(([.F350]+273.15)/100)+[.I350]*(-0.033096+0.014259*(([.F350]+273.15)/100)-0.0017*((([.F350]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G351]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J351]*112/[.J$499])*([.H351]/273)","v":0},{"$f":"of:=[.K351]/[.M351]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F351]+273.15))+143.3483*LN(([.F351]+273.15)/100)-21.8492*(([.F351]+273.15)/100)+[.I351]*(-0.033096+0.014259*(([.F351]+273.15)/100)-0.0017*((([.F351]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G352]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J352]*112/[.J$499])*([.H352]/273)","v":0},{"$f":"of:=[.K352]/[.M352]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F352]+273.15))+143.3483*LN(([.F352]+273.15)/100)-21.8492*(([.F352]+273.15)/100)+[.I352]*(-0.033096+0.014259*(([.F352]+273.15)/100)-0.0017*((([.F352]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G353]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J353]*112/[.J$499])*([.H353]/273)","v":0},{"$f":"of:=[.K353]/[.M353]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F353]+273.15))+143.3483*LN(([.F353]+273.15)/100)-21.8492*(([.F353]+273.15)/100)+[.I353]*(-0.033096+0.014259*(([.F353]+273.15)/100)-0.0017*((([.F353]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G354]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J354]*112/[.J$499])*([.H354]/273)","v":0},{"$f":"of:=[.K354]/[.M354]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F354]+273.15))+143.3483*LN(([.F354]+273.15)/100)-21.8492*(([.F354]+273.15)/100)+[.I354]*(-0.033096+0.014259*(([.F354]+273.15)/100)-0.0017*((([.F354]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G355]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J355]*112/[.J$499])*([.H355]/273)","v":0},{"$f":"of:=[.K355]/[.M355]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F355]+273.15))+143.3483*LN(([.F355]+273.15)/100)-21.8492*(([.F355]+273.15)/100)+[.I355]*(-0.033096+0.014259*(([.F355]+273.15)/100)-0.0017*((([.F355]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G356]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J356]*112/[.J$499])*([.H356]/273)","v":0},{"$f":"of:=[.K356]/[.M356]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F356]+273.15))+143.3483*LN(([.F356]+273.15)/100)-21.8492*(([.F356]+273.15)/100)+[.I356]*(-0.033096+0.014259*(([.F356]+273.15)/100)-0.0017*((([.F356]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G357]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J357]*112/[.J$499])*([.H357]/273)","v":0},{"$f":"of:=[.K357]/[.M357]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F357]+273.15))+143.3483*LN(([.F357]+273.15)/100)-21.8492*(([.F357]+273.15)/100)+[.I357]*(-0.033096+0.014259*(([.F357]+273.15)/100)-0.0017*((([.F357]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G360]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J360]*112/[.J$499])*([.H360]/273)","v":0},{"$f":"of:=[.K360]/[.M360]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F360]+273.15))+143.3483*LN(([.F360]+273.15)/100)-21.8492*(([.F360]+273.15)/100)+[.I360]*(-0.033096+0.014259*(([.F360]+273.15)/100)-0.0017*((([.F360]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G361]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J361]*112/[.J$499])*([.H361]/273)","v":0},{"$f":"of:=[.K361]/[.M361]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F361]+273.15))+143.3483*LN(([.F361]+273.15)/100)-21.8492*(([.F361]+273.15)/100)+[.I361]*(-0.033096+0.014259*(([.F361]+273.15)/100)-0.0017*((([.F361]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G362]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J362]*112/[.J$499])*([.H362]/273)","v":0},{"$f":"of:=[.K362]/[.M362]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F362]+273.15))+143.3483*LN(([.F362]+273.15)/100)-21.8492*(([.F362]+273.15)/100)+[.I362]*(-0.033096+0.014259*(([.F362]+273.15)/100)-0.0017*((([.F362]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G363]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J363]*112/[.J$499])*([.H363]/273)","v":0},{"$f":"of:=[.K363]/[.M363]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F363]+273.15))+143.3483*LN(([.F363]+273.15)/100)-21.8492*(([.F363]+273.15)/100)+[.I363]*(-0.033096+0.014259*(([.F363]+273.15)/100)-0.0017*((([.F363]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G364]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J364]*112/[.J$499])*([.H364]/273)","v":0},{"$f":"of:=[.K364]/[.M364]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F364]+273.15))+143.3483*LN(([.F364]+273.15)/100)-21.8492*(([.F364]+273.15)/100)+[.I364]*(-0.033096+0.014259*(([.F364]+273.15)/100)-0.0017*((([.F364]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G365]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J365]*112/[.J$499])*([.H365]/273)","v":0},{"$f":"of:=[.K365]/[.M365]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F365]+273.15))+143.3483*LN(([.F365]+273.15)/100)-21.8492*(([.F365]+273.15)/100)+[.I365]*(-0.033096+0.014259*(([.F365]+273.15)/100)-0.0017*((([.F365]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G366]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J366]*112/[.J$499])*([.H366]/273)","v":0},{"$f":"of:=[.K366]/[.M366]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F366]+273.15))+143.3483*LN(([.F366]+273.15)/100)-21.8492*(([.F366]+273.15)/100)+[.I366]*(-0.033096+0.014259*(([.F366]+273.15)/100)-0.0017*((([.F366]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G367]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J367]*112/[.J$499])*([.H367]/273)","v":0},{"$f":"of:=[.K367]/[.M367]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F367]+273.15))+143.3483*LN(([.F367]+273.15)/100)-21.8492*(([.F367]+273.15)/100)+[.I367]*(-0.033096+0.014259*(([.F367]+273.15)/100)-0.0017*((([.F367]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G368]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J368]*112/[.J$499])*([.H368]/273)","v":0},{"$f":"of:=[.K368]/[.M368]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F368]+273.15))+143.3483*LN(([.F368]+273.15)/100)-21.8492*(([.F368]+273.15)/100)+[.I368]*(-0.033096+0.014259*(([.F368]+273.15)/100)-0.0017*((([.F368]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G369]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J369]*112/[.J$499])*([.H369]/273)","v":0},{"$f":"of:=[.K369]/[.M369]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F369]+273.15))+143.3483*LN(([.F369]+273.15)/100)-21.8492*(([.F369]+273.15)/100)+[.I369]*(-0.033096+0.014259*(([.F369]+273.15)/100)-0.0017*((([.F369]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G372]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J372]*112/[.J$499])*([.H372]/273)","v":0},{"$f":"of:=[.K372]/[.M372]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F372]+273.15))+143.3483*LN(([.F372]+273.15)/100)-21.8492*(([.F372]+273.15)/100)+[.I372]*(-0.033096+0.014259*(([.F372]+273.15)/100)-0.0017*((([.F372]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G373]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J373]*112/[.J$499])*([.H373]/273)","v":0},{"$f":"of:=[.K373]/[.M373]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F373]+273.15))+143.3483*LN(([.F373]+273.15)/100)-21.8492*(([.F373]+273.15)/100)+[.I373]*(-0.033096+0.014259*(([.F373]+273.15)/100)-0.0017*((([.F373]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G374]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J374]*112/[.J$499])*([.H374]/273)","v":0},{"$f":"of:=[.K374]/[.M374]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F374]+273.15))+143.3483*LN(([.F374]+273.15)/100)-21.8492*(([.F374]+273.15)/100)+[.I374]*(-0.033096+0.014259*(([.F374]+273.15)/100)-0.0017*((([.F374]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G375]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J375]*112/[.J$499])*([.H375]/273)","v":0},{"$f":"of:=[.K375]/[.M375]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F375]+273.15))+143.3483*LN(([.F375]+273.15)/100)-21.8492*(([.F375]+273.15)/100)+[.I375]*(-0.033096+0.014259*(([.F375]+273.15)/100)-0.0017*((([.F375]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G376]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J376]*112/[.J$499])*([.H376]/273)","v":0},{"$f":"of:=[.K376]/[.M376]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F376]+273.15))+143.3483*LN(([.F376]+273.15)/100)-21.8492*(([.F376]+273.15)/100)+[.I376]*(-0.033096+0.014259*(([.F376]+273.15)/100)-0.0017*((([.F376]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G377]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J377]*112/[.J$499])*([.H377]/273)","v":0},{"$f":"of:=[.K377]/[.M377]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F377]+273.15))+143.3483*LN(([.F377]+273.15)/100)-21.8492*(([.F377]+273.15)/100)+[.I377]*(-0.033096+0.014259*(([.F377]+273.15)/100)-0.0017*((([.F377]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G378]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J378]*112/[.J$499])*([.H378]/273)","v":0},{"$f":"of:=[.K378]/[.M378]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F378]+273.15))+143.3483*LN(([.F378]+273.15)/100)-21.8492*(([.F378]+273.15)/100)+[.I378]*(-0.033096+0.014259*(([.F378]+273.15)/100)-0.0017*((([.F378]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G379]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J379]*112/[.J$499])*([.H379]/273)","v":0},{"$f":"of:=[.K379]/[.M379]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F379]+273.15))+143.3483*LN(([.F379]+273.15)/100)-21.8492*(([.F379]+273.15)/100)+[.I379]*(-0.033096+0.014259*(([.F379]+273.15)/100)-0.0017*((([.F379]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G380]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J380]*112/[.J$499])*([.H380]/273)","v":0},{"$f":"of:=[.K380]/[.M380]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F380]+273.15))+143.3483*LN(([.F380]+273.15)/100)-21.8492*(([.F380]+273.15)/100)+[.I380]*(-0.033096+0.014259*(([.F380]+273.15)/100)-0.0017*((([.F380]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G381]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J381]*112/[.J$499])*([.H381]/273)","v":0},{"$f":"of:=[.K381]/[.M381]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F381]+273.15))+143.3483*LN(([.F381]+273.15)/100)-21.8492*(([.F381]+273.15)/100)+[.I381]*(-0.033096+0.014259*(([.F381]+273.15)/100)-0.0017*((([.F381]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G384]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J384]*112/[.J$499])*([.H384]/273)","v":0},{"$f":"of:=[.K384]/[.M384]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F384]+273.15))+143.3483*LN(([.F384]+273.15)/100)-21.8492*(([.F384]+273.15)/100)+[.I384]*(-0.033096+0.014259*(([.F384]+273.15)/100)-0.0017*((([.F384]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G385]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J385]*112/[.J$499])*([.H385]/273)","v":0},{"$f":"of:=[.K385]/[.M385]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F385]+273.15))+143.3483*LN(([.F385]+273.15)/100)-21.8492*(([.F385]+273.15)/100)+[.I385]*(-0.033096+0.014259*(([.F385]+273.15)/100)-0.0017*((([.F385]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G386]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J386]*112/[.J$499])*([.H386]/273)","v":0},{"$f":"of:=[.K386]/[.M386]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F386]+273.15))+143.3483*LN(([.F386]+273.15)/100)-21.8492*(([.F386]+273.15)/100)+[.I386]*(-0.033096+0.014259*(([.F386]+273.15)/100)-0.0017*((([.F386]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G387]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J387]*112/[.J$499])*([.H387]/273)","v":0},{"$f":"of:=[.K387]/[.M387]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F387]+273.15))+143.3483*LN(([.F387]+273.15)/100)-21.8492*(([.F387]+273.15)/100)+[.I387]*(-0.033096+0.014259*(([.F387]+273.15)/100)-0.0017*((([.F387]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G388]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J388]*112/[.J$499])*([.H388]/273)","v":0},{"$f":"of:=[.K388]/[.M388]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F388]+273.15))+143.3483*LN(([.F388]+273.15)/100)-21.8492*(([.F388]+273.15)/100)+[.I388]*(-0.033096+0.014259*(([.F388]+273.15)/100)-0.0017*((([.F388]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G389]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J389]*112/[.J$499])*([.H389]/273)","v":0},{"$f":"of:=[.K389]/[.M389]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F389]+273.15))+143.3483*LN(([.F389]+273.15)/100)-21.8492*(([.F389]+273.15)/100)+[.I389]*(-0.033096+0.014259*(([.F389]+273.15)/100)-0.0017*((([.F389]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G390]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J390]*112/[.J$499])*([.H390]/273)","v":0},{"$f":"of:=[.K390]/[.M390]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F390]+273.15))+143.3483*LN(([.F390]+273.15)/100)-21.8492*(([.F390]+273.15)/100)+[.I390]*(-0.033096+0.014259*(([.F390]+273.15)/100)-0.0017*((([.F390]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G391]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J391]*112/[.J$499])*([.H391]/273)","v":0},{"$f":"of:=[.K391]/[.M391]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F391]+273.15))+143.3483*LN(([.F391]+273.15)/100)-21.8492*(([.F391]+273.15)/100)+[.I391]*(-0.033096+0.014259*(([.F391]+273.15)/100)-0.0017*((([.F391]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G392]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J392]*112/[.J$499])*([.H392]/273)","v":0},{"$f":"of:=[.K392]/[.M392]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F392]+273.15))+143.3483*LN(([.F392]+273.15)/100)-21.8492*(([.F392]+273.15)/100)+[.I392]*(-0.033096+0.014259*(([.F392]+273.15)/100)-0.0017*((([.F392]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G393]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J393]*112/[.J$499])*([.H393]/273)","v":0},{"$f":"of:=[.K393]/[.M393]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F393]+273.15))+143.3483*LN(([.F393]+273.15)/100)-21.8492*(([.F393]+273.15)/100)+[.I393]*(-0.033096+0.014259*(([.F393]+273.15)/100)-0.0017*((([.F393]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G396]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J396]*112/[.J$499])*([.H396]/273)","v":0},{"$f":"of:=[.K396]/[.M396]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F396]+273.15))+143.3483*LN(([.F396]+273.15)/100)-21.8492*(([.F396]+273.15)/100)+[.I396]*(-0.033096+0.014259*(([.F396]+273.15)/100)-0.0017*((([.F396]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G397]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J397]*112/[.J$499])*([.H397]/273)","v":0},{"$f":"of:=[.K397]/[.M397]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F397]+273.15))+143.3483*LN(([.F397]+273.15)/100)-21.8492*(([.F397]+273.15)/100)+[.I397]*(-0.033096+0.014259*(([.F397]+273.15)/100)-0.0017*((([.F397]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G398]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J398]*112/[.J$499])*([.H398]/273)","v":0},{"$f":"of:=[.K398]/[.M398]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F398]+273.15))+143.3483*LN(([.F398]+273.15)/100)-21.8492*(([.F398]+273.15)/100)+[.I398]*(-0.033096+0.014259*(([.F398]+273.15)/100)-0.0017*((([.F398]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G399]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J399]*112/[.J$499])*([.H399]/273)","v":0},{"$f":"of:=[.K399]/[.M399]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F399]+273.15))+143.3483*LN(([.F399]+273.15)/100)-21.8492*(([.F399]+273.15)/100)+[.I399]*(-0.033096+0.014259*(([.F399]+273.15)/100)-0.0017*((([.F399]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G400]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J400]*112/[.J$499])*([.H400]/273)","v":0},{"$f":"of:=[.K400]/[.M400]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F400]+273.15))+143.3483*LN(([.F400]+273.15)/100)-21.8492*(([.F400]+273.15)/100)+[.I400]*(-0.033096+0.014259*(([.F400]+273.15)/100)-0.0017*((([.F400]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G401]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J401]*112/[.J$499])*([.H401]/273)","v":0},{"$f":"of:=[.K401]/[.M401]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F401]+273.15))+143.3483*LN(([.F401]+273.15)/100)-21.8492*(([.F401]+273.15)/100)+[.I401]*(-0.033096+0.014259*(([.F401]+273.15)/100)-0.0017*((([.F401]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G402]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J402]*112/[.J$499])*([.H402]/273)","v":0},{"$f":"of:=[.K402]/[.M402]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F402]+273.15))+143.3483*LN(([.F402]+273.15)/100)-21.8492*(([.F402]+273.15)/100)+[.I402]*(-0.033096+0.014259*(([.F402]+273.15)/100)-0.0017*((([.F402]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G403]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J403]*112/[.J$499])*([.H403]/273)","v":0},{"$f":"of:=[.K403]/[.M403]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F403]+273.15))+143.3483*LN(([.F403]+273.15)/100)-21.8492*(([.F403]+273.15)/100)+[.I403]*(-0.033096+0.014259*(([.F403]+273.15)/100)-0.0017*((([.F403]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G404]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J404]*112/[.J$499])*([.H404]/273)","v":0},{"$f":"of:=[.K404]/[.M404]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F404]+273.15))+143.3483*LN(([.F404]+273.15)/100)-21.8492*(([.F404]+273.15)/100)+[.I404]*(-0.033096+0.014259*(([.F404]+273.15)/100)-0.0017*((([.F404]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G405]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J405]*112/[.J$499])*([.H405]/273)","v":0},{"$f":"of:=[.K405]/[.M405]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F405]+273.15))+143.3483*LN(([.F405]+273.15)/100)-21.8492*(([.F405]+273.15)/100)+[.I405]*(-0.033096+0.014259*(([.F405]+273.15)/100)-0.0017*((([.F405]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G408]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J408]*112/[.J$499])*([.H408]/273)","v":0},{"$f":"of:=[.K408]/[.M408]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F408]+273.15))+143.3483*LN(([.F408]+273.15)/100)-21.8492*(([.F408]+273.15)/100)+[.I408]*(-0.033096+0.014259*(([.F408]+273.15)/100)-0.0017*((([.F408]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G409]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J409]*112/[.J$499])*([.H409]/273)","v":0},{"$f":"of:=[.K409]/[.M409]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F409]+273.15))+143.3483*LN(([.F409]+273.15)/100)-21.8492*(([.F409]+273.15)/100)+[.I409]*(-0.033096+0.014259*(([.F409]+273.15)/100)-0.0017*((([.F409]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G410]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J410]*112/[.J$499])*([.H410]/273)","v":0},{"$f":"of:=[.K410]/[.M410]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F410]+273.15))+143.3483*LN(([.F410]+273.15)/100)-21.8492*(([.F410]+273.15)/100)+[.I410]*(-0.033096+0.014259*(([.F410]+273.15)/100)-0.0017*((([.F410]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G411]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J411]*112/[.J$499])*([.H411]/273)","v":0},{"$f":"of:=[.K411]/[.M411]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F411]+273.15))+143.3483*LN(([.F411]+273.15)/100)-21.8492*(([.F411]+273.15)/100)+[.I411]*(-0.033096+0.014259*(([.F411]+273.15)/100)-0.0017*((([.F411]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G412]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J412]*112/[.J$499])*([.H412]/273)","v":0},{"$f":"of:=[.K412]/[.M412]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F412]+273.15))+143.3483*LN(([.F412]+273.15)/100)-21.8492*(([.F412]+273.15)/100)+[.I412]*(-0.033096+0.014259*(([.F412]+273.15)/100)-0.0017*((([.F412]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G413]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J413]*112/[.J$499])*([.H413]/273)","v":0},{"$f":"of:=[.K413]/[.M413]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F413]+273.15))+143.3483*LN(([.F413]+273.15)/100)-21.8492*(([.F413]+273.15)/100)+[.I413]*(-0.033096+0.014259*(([.F413]+273.15)/100)-0.0017*((([.F413]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G414]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J414]*112/[.J$499])*([.H414]/273)","v":0},{"$f":"of:=[.K414]/[.M414]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F414]+273.15))+143.3483*LN(([.F414]+273.15)/100)-21.8492*(([.F414]+273.15)/100)+[.I414]*(-0.033096+0.014259*(([.F414]+273.15)/100)-0.0017*((([.F414]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G415]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J415]*112/[.J$499])*([.H415]/273)","v":0},{"$f":"of:=[.K415]/[.M415]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F415]+273.15))+143.3483*LN(([.F415]+273.15)/100)-21.8492*(([.F415]+273.15)/100)+[.I415]*(-0.033096+0.014259*(([.F415]+273.15)/100)-0.0017*((([.F415]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G416]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J416]*112/[.J$499])*([.H416]/273)","v":0},{"$f":"of:=[.K416]/[.M416]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F416]+273.15))+143.3483*LN(([.F416]+273.15)/100)-21.8492*(([.F416]+273.15)/100)+[.I416]*(-0.033096+0.014259*(([.F416]+273.15)/100)-0.0017*((([.F416]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G417]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J417]*112/[.J$499])*([.H417]/273)","v":0},{"$f":"of:=[.K417]/[.M417]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F417]+273.15))+143.3483*LN(([.F417]+273.15)/100)-21.8492*(([.F417]+273.15)/100)+[.I417]*(-0.033096+0.014259*(([.F417]+273.15)/100)-0.0017*((([.F417]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G420]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J420]*112/[.J$499])*([.H420]/273)","v":0},{"$f":"of:=[.K420]/[.M420]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F420]+273.15))+143.3483*LN(([.F420]+273.15)/100)-21.8492*(([.F420]+273.15)/100)+[.I420]*(-0.033096+0.014259*(([.F420]+273.15)/100)-0.0017*((([.F420]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G421]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J421]*112/[.J$499])*([.H421]/273)","v":0},{"$f":"of:=[.K421]/[.M421]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F421]+273.15))+143.3483*LN(([.F421]+273.15)/100)-21.8492*(([.F421]+273.15)/100)+[.I421]*(-0.033096+0.014259*(([.F421]+273.15)/100)-0.0017*((([.F421]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G422]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J422]*112/[.J$499])*([.H422]/273)","v":0},{"$f":"of:=[.K422]/[.M422]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F422]+273.15))+143.3483*LN(([.F422]+273.15)/100)-21.8492*(([.F422]+273.15)/100)+[.I422]*(-0.033096+0.014259*(([.F422]+273.15)/100)-0.0017*((([.F422]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G423]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J423]*112/[.J$499])*([.H423]/273)","v":0},{"$f":"of:=[.K423]/[.M423]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F423]+273.15))+143.3483*LN(([.F423]+273.15)/100)-21.8492*(([.F423]+273.15)/100)+[.I423]*(-0.033096+0.014259*(([.F423]+273.15)/100)-0.0017*((([.F423]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G424]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J424]*112/[.J$499])*([.H424]/273)","v":0},{"$f":"of:=[.K424]/[.M424]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F424]+273.15))+143.3483*LN(([.F424]+273.15)/100)-21.8492*(([.F424]+273.15)/100)+[.I424]*(-0.033096+0.014259*(([.F424]+273.15)/100)-0.0017*((([.F424]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G425]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J425]*112/[.J$499])*([.H425]/273)","v":0},{"$f":"of:=[.K425]/[.M425]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F425]+273.15))+143.3483*LN(([.F425]+273.15)/100)-21.8492*(([.F425]+273.15)/100)+[.I425]*(-0.033096+0.014259*(([.F425]+273.15)/100)-0.0017*((([.F425]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G426]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J426]*112/[.J$499])*([.H426]/273)","v":0},{"$f":"of:=[.K426]/[.M426]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F426]+273.15))+143.3483*LN(([.F426]+273.15)/100)-21.8492*(([.F426]+273.15)/100)+[.I426]*(-0.033096+0.014259*(([.F426]+273.15)/100)-0.0017*((([.F426]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G427]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J427]*112/[.J$499])*([.H427]/273)","v":0},{"$f":"of:=[.K427]/[.M427]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F427]+273.15))+143.3483*LN(([.F427]+273.15)/100)-21.8492*(([.F427]+273.15)/100)+[.I427]*(-0.033096+0.014259*(([.F427]+273.15)/100)-0.0017*((([.F427]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G428]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J428]*112/[.J$499])*([.H428]/273)","v":0},{"$f":"of:=[.K428]/[.M428]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F428]+273.15))+143.3483*LN(([.F428]+273.15)/100)-21.8492*(([.F428]+273.15)/100)+[.I428]*(-0.033096+0.014259*(([.F428]+273.15)/100)-0.0017*((([.F428]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G429]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J429]*112/[.J$499])*([.H429]/273)","v":0},{"$f":"of:=[.K429]/[.M429]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F429]+273.15))+143.3483*LN(([.F429]+273.15)/100)-21.8492*(([.F429]+273.15)/100)+[.I429]*(-0.033096+0.014259*(([.F429]+273.15)/100)-0.0017*((([.F429]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G432]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J432]*112/[.J$499])*([.H432]/273)","v":0},{"$f":"of:=[.K432]/[.M432]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F432]+273.15))+143.3483*LN(([.F432]+273.15)/100)-21.8492*(([.F432]+273.15)/100)+[.I432]*(-0.033096+0.014259*(([.F432]+273.15)/100)-0.0017*((([.F432]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G433]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J433]*112/[.J$499])*([.H433]/273)","v":0},{"$f":"of:=[.K433]/[.M433]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F433]+273.15))+143.3483*LN(([.F433]+273.15)/100)-21.8492*(([.F433]+273.15)/100)+[.I433]*(-0.033096+0.014259*(([.F433]+273.15)/100)-0.0017*((([.F433]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G434]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J434]*112/[.J$499])*([.H434]/273)","v":0},{"$f":"of:=[.K434]/[.M434]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F434]+273.15))+143.3483*LN(([.F434]+273.15)/100)-21.8492*(([.F434]+273.15)/100)+[.I434]*(-0.033096+0.014259*(([.F434]+273.15)/100)-0.0017*((([.F434]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G435]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J435]*112/[.J$499])*([.H435]/273)","v":0},{"$f":"of:=[.K435]/[.M435]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F435]+273.15))+143.3483*LN(([.F435]+273.15)/100)-21.8492*(([.F435]+273.15)/100)+[.I435]*(-0.033096+0.014259*(([.F435]+273.15)/100)-0.0017*((([.F435]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G436]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J436]*112/[.J$499])*([.H436]/273)","v":0},{"$f":"of:=[.K436]/[.M436]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F436]+273.15))+143.3483*LN(([.F436]+273.15)/100)-21.8492*(([.F436]+273.15)/100)+[.I436]*(-0.033096+0.014259*(([.F436]+273.15)/100)-0.0017*((([.F436]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G437]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J437]*112/[.J$499])*([.H437]/273)","v":0},{"$f":"of:=[.K437]/[.M437]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F437]+273.15))+143.3483*LN(([.F437]+273.15)/100)-21.8492*(([.F437]+273.15)/100)+[.I437]*(-0.033096+0.014259*(([.F437]+273.15)/100)-0.0017*((([.F437]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G438]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J438]*112/[.J$499])*([.H438]/273)","v":0},{"$f":"of:=[.K438]/[.M438]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F438]+273.15))+143.3483*LN(([.F438]+273.15)/100)-21.8492*(([.F438]+273.15)/100)+[.I438]*(-0.033096+0.014259*(([.F438]+273.15)/100)-0.0017*((([.F438]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G439]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J439]*112/[.J$499])*([.H439]/273)","v":0},{"$f":"of:=[.K439]/[.M439]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F439]+273.15))+143.3483*LN(([.F439]+273.15)/100)-21.8492*(([.F439]+273.15)/100)+[.I439]*(-0.033096+0.014259*(([.F439]+273.15)/100)-0.0017*((([.F439]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G440]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J440]*112/[.J$499])*([.H440]/273)","v":0},{"$f":"of:=[.K440]/[.M440]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F440]+273.15))+143.3483*LN(([.F440]+273.15)/100)-21.8492*(([.F440]+273.15)/100)+[.I440]*(-0.033096+0.014259*(([.F440]+273.15)/100)-0.0017*((([.F440]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G441]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J441]*112/[.J$499])*([.H441]/273)","v":0},{"$f":"of:=[.K441]/[.M441]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F441]+273.15))+143.3483*LN(([.F441]+273.15)/100)-21.8492*(([.F441]+273.15)/100)+[.I441]*(-0.033096+0.014259*(([.F441]+273.15)/100)-0.0017*((([.F441]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G444]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J444]*112/[.J$499])*([.H444]/273)","v":0},{"$f":"of:=[.K444]/[.M444]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F444]+273.15))+143.3483*LN(([.F444]+273.15)/100)-21.8492*(([.F444]+273.15)/100)+[.I444]*(-0.033096+0.014259*(([.F444]+273.15)/100)-0.0017*((([.F444]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G445]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J445]*112/[.J$499])*([.H445]/273)","v":0},{"$f":"of:=[.K445]/[.M445]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F445]+273.15))+143.3483*LN(([.F445]+273.15)/100)-21.8492*(([.F445]+273.15)/100)+[.I445]*(-0.033096+0.014259*(([.F445]+273.15)/100)-0.0017*((([.F445]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G446]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J446]*112/[.J$499])*([.H446]/273)","v":0},{"$f":"of:=[.K446]/[.M446]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F446]+273.15))+143.3483*LN(([.F446]+273.15)/100)-21.8492*(([.F446]+273.15)/100)+[.I446]*(-0.033096+0.014259*(([.F446]+273.15)/100)-0.0017*((([.F446]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G447]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J447]*112/[.J$499])*([.H447]/273)","v":0},{"$f":"of:=[.K447]/[.M447]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F447]+273.15))+143.3483*LN(([.F447]+273.15)/100)-21.8492*(([.F447]+273.15)/100)+[.I447]*(-0.033096+0.014259*(([.F447]+273.15)/100)-0.0017*((([.F447]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G448]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J448]*112/[.J$499])*([.H448]/273)","v":0},{"$f":"of:=[.K448]/[.M448]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F448]+273.15))+143.3483*LN(([.F448]+273.15)/100)-21.8492*(([.F448]+273.15)/100)+[.I448]*(-0.033096+0.014259*(([.F448]+273.15)/100)-0.0017*((([.F448]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G449]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J449]*112/[.J$499])*([.H449]/273)","v":0},{"$f":"of:=[.K449]/[.M449]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F449]+273.15))+143.3483*LN(([.F449]+273.15)/100)-21.8492*(([.F449]+273.15)/100)+[.I449]*(-0.033096+0.014259*(([.F449]+273.15)/100)-0.0017*((([.F449]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G450]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J450]*112/[.J$499])*([.H450]/273)","v":0},{"$f":"of:=[.K450]/[.M450]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F450]+273.15))+143.3483*LN(([.F450]+273.15)/100)-21.8492*(([.F450]+273.15)/100)+[.I450]*(-0.033096+0.014259*(([.F450]+273.15)/100)-0.0017*((([.F450]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G451]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J451]*112/[.J$499])*([.H451]/273)","v":0},{"$f":"of:=[.K451]/[.M451]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F451]+273.15))+143.3483*LN(([.F451]+273.15)/100)-21.8492*(([.F451]+273.15)/100)+[.I451]*(-0.033096+0.014259*(([.F451]+273.15)/100)-0.0017*((([.F451]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G452]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J452]*112/[.J$499])*([.H452]/273)","v":0},{"$f":"of:=[.K452]/[.M452]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F452]+273.15))+143.3483*LN(([.F452]+273.15)/100)-21.8492*(([.F452]+273.15)/100)+[.I452]*(-0.033096+0.014259*(([.F452]+273.15)/100)-0.0017*((([.F452]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G453]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J453]*112/[.J$499])*([.H453]/273)","v":0},{"$f":"of:=[.K453]/[.M453]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F453]+273.15))+143.3483*LN(([.F453]+273.15)/100)-21.8492*(([.F453]+273.15)/100)+[.I453]*(-0.033096+0.014259*(([.F453]+273.15)/100)-0.0017*((([.F453]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G456]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J456]*112/[.J$499])*([.H456]/273)","v":0},{"$f":"of:=[.K456]/[.M456]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F456]+273.15))+143.3483*LN(([.F456]+273.15)/100)-21.8492*(([.F456]+273.15)/100)+[.I456]*(-0.033096+0.014259*(([.F456]+273.15)/100)-0.0017*((([.F456]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G457]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J457]*112/[.J$499])*([.H457]/273)","v":0},{"$f":"of:=[.K457]/[.M457]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F457]+273.15))+143.3483*LN(([.F457]+273.15)/100)-21.8492*(([.F457]+273.15)/100)+[.I457]*(-0.033096+0.014259*(([.F457]+273.15)/100)-0.0017*((([.F457]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G458]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J458]*112/[.J$499])*([.H458]/273)","v":0},{"$f":"of:=[.K458]/[.M458]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F458]+273.15))+143.3483*LN(([.F458]+273.15)/100)-21.8492*(([.F458]+273.15)/100)+[.I458]*(-0.033096+0.014259*(([.F458]+273.15)/100)-0.0017*((([.F458]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G459]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J459]*112/[.J$499])*([.H459]/273)","v":0},{"$f":"of:=[.K459]/[.M459]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F459]+273.15))+143.3483*LN(([.F459]+273.15)/100)-21.8492*(([.F459]+273.15)/100)+[.I459]*(-0.033096+0.014259*(([.F459]+273.15)/100)-0.0017*((([.F459]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G460]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J460]*112/[.J$499])*([.H460]/273)","v":0},{"$f":"of:=[.K460]/[.M460]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F460]+273.15))+143.3483*LN(([.F460]+273.15)/100)-21.8492*(([.F460]+273.15)/100)+[.I460]*(-0.033096+0.014259*(([.F460]+273.15)/100)-0.0017*((([.F460]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G461]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J461]*112/[.J$499])*([.H461]/273)","v":0},{"$f":"of:=[.K461]/[.M461]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F461]+273.15))+143.3483*LN(([.F461]+273.15)/100)-21.8492*(([.F461]+273.15)/100)+[.I461]*(-0.033096+0.014259*(([.F461]+273.15)/100)-0.0017*((([.F461]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G462]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J462]*112/[.J$499])*([.H462]/273)","v":0},{"$f":"of:=[.K462]/[.M462]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F462]+273.15))+143.3483*LN(([.F462]+273.15)/100)-21.8492*(([.F462]+273.15)/100)+[.I462]*(-0.033096+0.014259*(([.F462]+273.15)/100)-0.0017*((([.F462]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G463]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J463]*112/[.J$499])*([.H463]/273)","v":0},{"$f":"of:=[.K463]/[.M463]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F463]+273.15))+143.3483*LN(([.F463]+273.15)/100)-21.8492*(([.F463]+273.15)/100)+[.I463]*(-0.033096+0.014259*(([.F463]+273.15)/100)-0.0017*((([.F463]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G464]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J464]*112/[.J$499])*([.H464]/273)","v":0},{"$f":"of:=[.K464]/[.M464]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F464]+273.15))+143.3483*LN(([.F464]+273.15)/100)-21.8492*(([.F464]+273.15)/100)+[.I464]*(-0.033096+0.014259*(([.F464]+273.15)/100)-0.0017*((([.F464]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G465]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J465]*112/[.J$499])*([.H465]/273)","v":0},{"$f":"of:=[.K465]/[.M465]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F465]+273.15))+143.3483*LN(([.F465]+273.15)/100)-21.8492*(([.F465]+273.15)/100)+[.I465]*(-0.033096+0.014259*(([.F465]+273.15)/100)-0.0017*((([.F465]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G468]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J468]*112/[.J$499])*([.H468]/273)","v":0},{"$f":"of:=[.K468]/[.M468]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F468]+273.15))+143.3483*LN(([.F468]+273.15)/100)-21.8492*(([.F468]+273.15)/100)+[.I468]*(-0.033096+0.014259*(([.F468]+273.15)/100)-0.0017*((([.F468]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G469]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J469]*112/[.J$499])*([.H469]/273)","v":0},{"$f":"of:=[.K469]/[.M469]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F469]+273.15))+143.3483*LN(([.F469]+273.15)/100)-21.8492*(([.F469]+273.15)/100)+[.I469]*(-0.033096+0.014259*(([.F469]+273.15)/100)-0.0017*((([.F469]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G470]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J470]*112/[.J$499])*([.H470]/273)","v":0},{"$f":"of:=[.K470]/[.M470]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F470]+273.15))+143.3483*LN(([.F470]+273.15)/100)-21.8492*(([.F470]+273.15)/100)+[.I470]*(-0.033096+0.014259*(([.F470]+273.15)/100)-0.0017*((([.F470]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G471]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J471]*112/[.J$499])*([.H471]/273)","v":0},{"$f":"of:=[.K471]/[.M471]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F471]+273.15))+143.3483*LN(([.F471]+273.15)/100)-21.8492*(([.F471]+273.15)/100)+[.I471]*(-0.033096+0.014259*(([.F471]+273.15)/100)-0.0017*((([.F471]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G472]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J472]*112/[.J$499])*([.H472]/273)","v":0},{"$f":"of:=[.K472]/[.M472]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F472]+273.15))+143.3483*LN(([.F472]+273.15)/100)-21.8492*(([.F472]+273.15)/100)+[.I472]*(-0.033096+0.014259*(([.F472]+273.15)/100)-0.0017*((([.F472]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G473]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J473]*112/[.J$499])*([.H473]/273)","v":0},{"$f":"of:=[.K473]/[.M473]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F473]+273.15))+143.3483*LN(([.F473]+273.15)/100)-21.8492*(([.F473]+273.15)/100)+[.I473]*(-0.033096+0.014259*(([.F473]+273.15)/100)-0.0017*((([.F473]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G474]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J474]*112/[.J$499])*([.H474]/273)","v":0},{"$f":"of:=[.K474]/[.M474]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F474]+273.15))+143.3483*LN(([.F474]+273.15)/100)-21.8492*(([.F474]+273.15)/100)+[.I474]*(-0.033096+0.014259*(([.F474]+273.15)/100)-0.0017*((([.F474]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G475]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J475]*112/[.J$499])*([.H475]/273)","v":0},{"$f":"of:=[.K475]/[.M475]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F475]+273.15))+143.3483*LN(([.F475]+273.15)/100)-21.8492*(([.F475]+273.15)/100)+[.I475]*(-0.033096+0.014259*(([.F475]+273.15)/100)-0.0017*((([.F475]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G476]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J476]*112/[.J$499])*([.H476]/273)","v":0},{"$f":"of:=[.K476]/[.M476]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F476]+273.15))+143.3483*LN(([.F476]+273.15)/100)-21.8492*(([.F476]+273.15)/100)+[.I476]*(-0.033096+0.014259*(([.F476]+273.15)/100)-0.0017*((([.F476]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G477]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J477]*112/[.J$499])*([.H477]/273)","v":0},{"$f":"of:=[.K477]/[.M477]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F477]+273.15))+143.3483*LN(([.F477]+273.15)/100)-21.8492*(([.F477]+273.15)/100)+[.I477]*(-0.033096+0.014259*(([.F477]+273.15)/100)-0.0017*((([.F477]+273.15)/100))^2))","v":10.218032355861}]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[{"$nulls":7},{"$f":"of:=[.G480]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J480]*112/[.J$499])*([.H480]/273)","v":0},{"$f":"of:=[.K480]/[.M480]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F480]+273.15))+143.3483*LN(([.F480]+273.15)/100)-21.8492*(([.F480]+273.15)/100)+[.I480]*(-0.033096+0.014259*(([.F480]+273.15)/100)-0.0017*((([.F480]+273.15)/100))^2))","v":10.218032355861}]
["POSIÇÃO GEOGRÁFICA",{"$nulls":6},{"$f":"of:=[.G481]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J481]*112/[.J$499])*([.H481]/273)","v":0},{"$f":"of:=[.K481]/[.M481]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F481]+273.15))+143.3483*LN(([.F481]+273.15)/100)-21.8492*(([.F481]+273.15)/100)+[.I481]*(-0.033096+0.014259*(([.F481]+273.15)/100)-0.0017*((([.F481]+273.15)/100))^2))","v":10.218032355861}]
["Latitude",{"$nulls":6},{"$f":"of:=[.G482]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J482]*112/[.J$499])*([.H482]/273)","v":0},{"$f":"of:=[.K482]/[.M482]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F482]+273.15))+143.3483*LN(([.F482]+273.15)/100)-21.8492*(([.F482]+273.15)/100)+[.I482]*(-0.033096+0.014259*(([.F482]+273.15)/100)-0.0017*((([.F482]+273.15)/100))^2))","v":10.218032355861}]
["Longitude",{"$nulls":6},{"$f":"of:=[.G483]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J483]*112/[.J$499])*([.H483]/273)","v":0},{"$f":"of:=[.K483]/[.M483]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F483]+273.15))+143.3483*LN(([.F483]+273.15)/100)-21.8492*(([.F483]+273.15)/100)+[.I483]*(-0.033096+0.014259*(([.F483]+273.15)/100)-0.0017*((([.F483]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G484]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J484]*112/[.J$499])*([.H484]/273)","v":0},{"$f":"of:=[.K484]/[.M484]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F484]+273.15))+143.3483*LN(([.F484]+273.15)/100)-21.8492*(([.F484]+273.15)/100)+[.I484]*(-0.033096+0.014259*(([.F484]+273.15)/100)-0.0017*((([.F484]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G485]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J485]*112/[.J$499])*([.H485]/273)","v":0},{"$f":"of:=[.K485]/[.M485]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F485]+273.15))+143.3483*LN(([.F485]+273.15)/100)-21.8492*(([.F485]+273.15)/100)+[.I485]*(-0.033096+0.014259*(([.F485]+273.15)/100)-0.0017*((([.F485]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G486]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J486]*112/[.J$499])*([.H486]/273)","v":0},{"$f":"of:=[.K486]/[.M486]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F486]+273.15))+143.3483*LN(([.F486]+273.15)/100)-21.8492*(([.F486]+273.15)/100)+[.I486]*(-0.033096+0.014259*(([.F486]+273.15)/100)-0.0017*((([.F486]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G487]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J487]*112/[.J$499])*([.H487]/273)","v":0},{"$f":"of:=[.K487]/[.M487]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F487]+273.15))+143.3483*LN(([.F487]+273.15)/100)-21.8492*(([.F487]+273.15)/100)+[.I487]*(-0.033096+0.014259*(([.F487]+273.15)/100)-0.0017*((([.F487]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G488]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J488]*112/[.J$499])*([.H488]/273)","v":0},{"$f":"of:=[.K488]/[.M488]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F488]+273.15))+143.3483*LN(([.F488]+273.15)/100)-21.8492*(([.F488]+273.15)/100)+[.I488]*(-0.033096+0.014259*(([.F488]+273.15)/100)-0.0017*((([.F488]+273.15)/100))^2))","v":10.218032355861}]
[{"$nulls":7},{"$f":"of:=[.G489]+273","v":273},null,null,{"$f":"of:=([.J$498]*[.J489]*112/[.J$499])*([.H489]/273)","v":0},{"$f":"of:=[.K489]/[.M489]","v":{"v":0,"d":"0%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F489]+273.15))+143.3483*LN(([.F489]+273.15)/100)-21.8492*(([.F489]+273.15)/100)+[.I489]*(-0.033096+0.014259*(([.F489]+273.15)/100)-0.0017*((([.F489]+273.15)/100))^2))","v":10.218032355861}]
[]
[]
[]
[]
[]
[]
[null,"DADOS COMPLEMENTARES"]
[]
[null,"FATOR DE CORREÇÃO TIOSSULFATO DE SÓDIO"]
[null,"VOLUME TOTAL DA ALICOTA UTILIZADA PARA ANÁLISE (ml)",{"$nulls":7},50]
[]
[null,"OBSERVAÇÕES GERAIS"]
[null," PERFIL Nº"]
[null,"Nº ESTAÇÕES "]
[null,"REALIZADAS"]
[null,null,"NÃO REALIZADAS"]
[null,"LEITURA DAS AMOSTRAS"]
[]
[null,"FEITO POR : "]
[null,"CONFERIDO POR: "]

{"sheet":"PIRXIII","index":1,"n_rows":145,"n_cols":15,"rows_kept":145,"merged":[[0,0,0,14],[1,0,1,14],[3,0,3,14],[5,0,5,14],[8,0,8,2],[8,3,8,12],[12,0,12,1],[13,1,13,2],[14,1,14,2],[20,0,20,1],[21,1,21,2],[22,1,22,2],[28,0,28,1],[29,1,29,2],[30,1,30,2],[36,0,36,1],[37,1,37,2],[38,1,38,2],[44,0,44,1],[45,1,45,2],[46,1,46,2],[52,0,52,1],[53,1,53,2],[54,1,54,2],[60,0,60,1],[61,1,61,2],[62,1,62,2],[68,0,68,1],[69,1,69,2],[70,1,70,2],[76,0,76,1],[77,1,77,2],[78,1,78,2],[84,0,84,1],[85,1,85,2],[86,1,86,2],[92,0,92,1],[93,1,93,2],[94,1,94,2],[100,0,100,1],[101,1,101,2],[102,1,102,2],[108,0,108,1],[109,1,109,2],[110,1,110,2],[116,0,116,1],[117,1,117,2],[118,1,118,2],[124,0,124,1],[125,1,125,2],[126,1,126,2],[132,1,133,9],[134,1,134,8],[135,1,135,8],[137,1,137,4],[138,1,138,3],[139,1,139,3],[140,1,140,3],[141,1,141,3],[143,3,143,5],[144,3,144,5]]}
["MARINHA DO BRASIL"]
["Noc ANTARES"]
[]
["CALCULO DO OXIGÊNIO DISSOLVIDO NA ÁGUA DO MAR"]
[]
["PIRATA XIII / 2011."]
[]
[]
["PERÍODO :",null,null,"11 de agosto à 05 de outubro."]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[1,"2011-08-24","PT08H01M00S",3150,2,27.0375,25.2,{"$f":"of:=[.G12]+273","v":298.2},34.9647,1.83,{"$f":"of:=([.J$135]*[.J12]*112/[.J$136])*([.H12]/273)","v":4.23212445442855},{"$f":"of:=[.K12]/[.M12]","v":{"v":0.92642521374794,"d":"93%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F12]+273.15))+143.3483*LN(([.F12]+273.15)/100)-21.8492*(([.F12]+273.15)/100)+[.I12]*(-0.033096+0.014259*(([.F12]+273.15)/100)-0.0017*((([.F12]+273.15)/100))^2))","v":4.56823107966492},null,4.3296]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},79,27.0156,25.6,{"$f":"of:=[.G13]+273","v":298.6},34.4739,1.82,{"$f":"of:=([.J$135]*[.J13]*112/[.J$136])*([.H13]/273)","v":4.2146439639808},{"$f":"of:=[.K13]/[.M13]","v":{"v":0.919726995130756,"d":"92%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F13]+273.15))+143.3483*LN(([.F13]+273.15)/100)-21.8492*(([.F13]+273.15)/100)+[.I13]*(-0.033096+0.014259*(([.F13]+273.15)/100)-0.0017*((([.F13]+273.15)/100))^2))","v":4.58249457316582},null,4.3353]
["Latitude","01º59,32 S",null,null,99,26.0305,25.2,{"$f":"of:=[.G14]+273","v":298.2},35.8505,1.83,{"$f":"of:=([.J$135]*[.J14]*112/[.J$136])*([.H14]/273)","v":4.23212445442855},{"$f":"of:=[.K14]/[.M14]","v":{"v":0.91555389919609,"d":"92%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F14]+273.15))+143.3483*LN(([.F14]+273.15)/100)-21.8492*(([.F14]+273.15)/100)+[.I14]*(-0.033096+0.014259*(([.F14]+273.15)/100)-0.0017*((([.F14]+273.15)/100))^2))","v":4.62247439298179},null,4.4961]
["Longitude","038º02,41 W",null,null,150,17.3745,26,{"$f":"of:=[.G15]+273","v":299},36.7684,1.76,{"$f":"of:=([.J$135]*[.J15]*112/[.J$136])*([.H15]/273)","v":4.081159402496},{"$f":"of:=[.K15]/[.M15]","v":{"v":0.759885798623276,"d":"76%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F15]+273.15))+143.3483*LN(([.F15]+273.15)/100)-21.8492*(([.F15]+273.15)/100)+[.I15]*(-0.033096+0.014259*(([.F15]+273.15)/100)-0.0017*((([.F15]+273.15)/100))^2))","v":5.37075361835956},null,4.0513]
[{"$nulls":4},598,5.9027,13.8,{"$f":"of:=[.G16]+273","v":286.8},36.44,1.73,{"$f":"of:=([.J$135]*[.J16]*112/[.J$136])*([.H16]/273)","v":3.84791040928098},{"$f":"of:=[.K16]/[.M16]","v":{"v":0.559875455429287,"d":"56%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F16]+273.15))+143.3483*LN(([.F16]+273.15)/100)-21.8492*(([.F16]+273.15)/100)+[.I16]*(-0.033096+0.014259*(([.F16]+273.15)/100)-0.0017*((([.F16]+273.15)/100))^2))","v":6.8727971050822},null,3.7717]
[{"$nulls":4},1496,4.3548,11.2,{"$f":"of:=[.G17]+273","v":284.2},36.3247,2.16,{"$f":"of:=([.J$135]*[.J17]*112/[.J$136])*([.H17]/273)","v":4.76077357237957},{"$f":"of:=[.K17]/[.M17]","v":{"v":0.666946442613016,"d":"67%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F17]+273.15))+143.3483*LN(([.F17]+273.15)/100)-21.8492*(([.F17]+273.15)/100)+[.I17]*(-0.033096+0.014259*(([.F17]+273.15)/100)-0.0017*((([.F17]+273.15)/100))^2))","v":7.13816472838123},null,5.0314]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[3,"2011-08-24","PT23H25M00S",4443,3,26.8992,25,{"$f":"of:=[.F20]+273","v":299.8992},36.0127,1.93,{"$f":"of:=([.J$135]*[.J20]*112/[.J$136])*([.H20]/273)","v":4.48882131531314},{"$f":"of:=[.K20]/[.M20]","v":{"v":0.986153457892157,"d":"99%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F20]+273.15))+143.3483*LN(([.F20]+273.15)/100)-21.8492*(([.F20]+273.15)/100)+[.I20]*(-0.033096+0.014259*(([.F20]+273.15)/100)-0.0017*((([.F20]+273.15)/100))^2))","v":4.55184867972548},null,4.3295]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},108,21.8556,22.2,{"$f":"of:=[.F21]+273","v":294.8556},36.4947,1.84,{"$f":"of:=([.J$135]*[.J21]*112/[.J$136])*([.H21]/273)","v":4.20752693817692},{"$f":"of:=[.K21]/[.M21]","v":{"v":0.849886429885201,"d":"85%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F21]+273.15))+143.3483*LN(([.F21]+273.15)/100)-21.8492*(([.F21]+273.15)/100)+[.I21]*(-0.033096+0.014259*(([.F21]+273.15)/100)-0.0017*((([.F21]+273.15)/100))^2))","v":4.95069316349157},null,4.0266]
["Latitude","00º00,24 S",null,null,128,18.6044,21.9,{"$f":"of:=[.G22]+273","v":294.9},35.9848,1.6,{"$f":"of:=([.J$135]*[.J22]*112/[.J$136])*([.H22]/273)","v":3.65927001458215},{"$f":"of:=[.K22]/[.M22]","v":{"v":0.694215125553437,"d":"69%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F22]+273.15))+143.3483*LN(([.F22]+273.15)/100)-21.8492*(([.F22]+273.15)/100)+[.I22]*(-0.033096+0.014259*(([.F22]+273.15)/100)-0.0017*((([.F22]+273.15)/100))^2))","v":5.27108943595105},null,3.7935]
["Longitude","038º00,34 W",null,null,400,8.4633,19.6,{"$f":"of:=[.G23]+273","v":292.6},34.7384,1.32,{"$f":"of:=([.J$135]*[.J23]*112/[.J$136])*([.H23]/273)","v":2.99535261163126},{"$f":"of:=[.K23]/[.M23]","v":{"v":0.457331045270302,"d":"46%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F23]+273.15))+143.3483*LN(([.F23]+273.15)/100)-21.8492*(([.F23]+273.15)/100)+[.I23]*(-0.033096+0.014259*(([.F23]+273.15)/100)-0.0017*((([.F23]+273.15)/100))^2))","v":6.54963760411428},null,2.9854]
[{"$nulls":4},685,5.4267,19.4,{"$f":"of:=[.G24]+273","v":292.4},34.5045,1.48,{"$f":"of:=([.J$135]*[.J24]*112/[.J$136])*([.H24]/273)","v":3.35613008085465},{"$f":"of:=[.K24]/[.M24]","v":{"v":0.476736067973891,"d":"48%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F24]+273.15))+143.3483*LN(([.F24]+273.15)/100)-21.8492*(([.F24]+273.15)/100)+[.I24]*(-0.033096+0.014259*(([.F24]+273.15)/100)-0.0017*((([.F24]+273.15)/100))^2))","v":7.03980736158283},null,3.3149]
[{"$nulls":4},1493,4.3067,19.3,{"$f":"of:=[.G25]+273","v":292.3},34.9731,2.21,{"$f":"of:=([.J$135]*[.J25]*112/[.J$136])*([.H25]/273)","v":5.0098046410432},{"$f":"of:=[.K25]/[.M25]","v":{"v":0.694767329056788,"d":"69%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F25]+273.15))+143.3483*LN(([.F25]+273.15)/100)-21.8492*(([.F25]+273.15)/100)+[.I25]*(-0.033096+0.014259*(([.F25]+273.15)/100)-0.0017*((([.F25]+273.15)/100))^2))","v":7.21076600974384},null,5.1594]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[5,"2011-08-25","PT15H42M00S",4447,3,27.8959,27.6,{"$f":"of:=[.G28]+273","v":300.6},35.6129,1.85,{"$f":"of:=([.J$135]*[.J28]*112/[.J$136])*([.H28]/273)","v":4.31281079986708},{"$f":"of:=[.K28]/[.M28]","v":{"v":0.960995032635875,"d":"96%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F28]+273.15))+143.3483*LN(([.F28]+273.15)/100)-21.8492*(([.F28]+273.15)/100)+[.I28]*(-0.033096+0.014259*(([.F28]+273.15)/100)-0.0017*((([.F28]+273.15)/100))^2))","v":4.4878596177939},null,4.3225]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},96,26.6107,27.4,{"$f":"of:=[.G29]+273","v":300.4},35.7691,1.68,{"$f":"of:=([.J$135]*[.J29]*112/[.J$136])*([.H29]/273)","v":3.91389266869957},{"$f":"of:=[.K29]/[.M29]","v":{"v":0.854567014200098,"d":"85%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F29]+273.15))+143.3483*LN(([.F29]+273.15)/100)-21.8492*(([.F29]+273.15)/100)+[.I29]*(-0.033096+0.014259*(([.F29]+273.15)/100)-0.0017*((([.F29]+273.15)/100))^2))","v":4.57997161564105},null,3.9147]
["Latitude","01º59,98 N",null,null,113,17.0566,22.9,{"$f":"of:=[.G30]+273","v":295.9},35.6961,0.89,{"$f":"of:=([.J$135]*[.J30]*112/[.J$136])*([.H30]/273)","v":2.04237118008271},{"$f":"of:=[.K30]/[.M30]","v":{"v":0.375515759500793,"d":"38%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F30]+273.15))+143.3483*LN(([.F30]+273.15)/100)-21.8492*(([.F30]+273.15)/100)+[.I30]*(-0.033096+0.014259*(([.F30]+273.15)/100)-0.0017*((([.F30]+273.15)/100))^2))","v":5.43884278731156},null,2.058]
["Longitude","038º00,34 W",null,null,495,7.1746,15.9,{"$f":"of:=[.G31]+273","v":288.9},34.6334,1.26,{"$f":"of:=([.J$135]*[.J31]*112/[.J$136])*([.H31]/273)","v":2.82304492007483},{"$f":"of:=[.K31]/[.M31]","v":{"v":0.418212999504996,"d":"42%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F31]+273.15))+143.3483*LN(([.F31]+273.15)/100)-21.8492*(([.F31]+273.15)/100)+[.I31]*(-0.033096+0.014259*(([.F31]+273.15)/100)-0.0017*((([.F31]+273.15)/100))^2))","v":6.75025626514775},null,2.8791]
[{"$nulls":4},698,5.798,15.1,{"$f":"of:=[.G32]+273","v":288.1},34.5175,1.5,{"$f":"of:=([.J$135]*[.J32]*112/[.J$136])*([.H32]/273)","v":3.35146137843692},{"$f":"of:=[.K32]/[.M32]","v":{"v":0.480349324522289,"d":"48%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F32]+273.15))+143.3483*LN(([.F32]+273.15)/100)-21.8492*(([.F32]+273.15)/100)+[.I32]*(-0.033096+0.014259*(([.F32]+273.15)/100)-0.0017*((([.F32]+273.15)/100))^2))","v":6.97713353041555},null,3.322]
[{"$nulls":4},1470,4.4042,14.5,{"$f":"of:=[.G33]+273","v":287.5},34.9722,2.24,{"$f":"of:=([.J$135]*[.J33]*112/[.J$136])*([.H33]/273)","v":4.99442584221539},{"$f":"of:=[.K33]/[.M33]","v":{"v":0.694280550587806,"d":"69%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F33]+273.15))+143.3483*LN(([.F33]+273.15)/100)-21.8492*(([.F33]+273.15)/100)+[.I33]*(-0.033096+0.014259*(([.F33]+273.15)/100)-0.0017*((([.F33]+273.15)/100))^2))","v":7.19367097059957},null,5.0361]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[7,"2011-08-26","PT06H15M00S",4085,3,28.3323,26.4,{"$f":"of:=[.G36]+273","v":299.4},35.6245,1.84,{"$f":"of:=([.J$135]*[.J36]*112/[.J$136])*([.H36]/273)","v":4.27237456331225},{"$f":"of:=[.K36]/[.M36]","v":{"v":0.792403886260823,"d":"79%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F36]+273.15))+143.3483*LN(([.F36]+273.15)/100)-21.8492*(([.F36]+273.15)/100)+[.J38]*(-0.033096+0.014259*(([.F36]+273.15)/100)-0.0017*((([.F36]+273.15)/100))^2))","v":5.39166280906651},null,4.2945]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},89,27.638,26.6,{"$f":"of:=[.G37]+273","v":299.6},35.9943,1.75,{"$f":"of:=([.J$135]*[.J37]*112/[.J$136])*([.H37]/273)","v":4.06611408241231},{"$f":"of:=[.K37]/[.M37]","v":{"v":0.745105895844937,"d":"75%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F37]+273.15))+143.3483*LN(([.F37]+273.15)/100)-21.8492*(([.F37]+273.15)/100)+[.J39]*(-0.033096+0.014259*(([.F37]+273.15)/100)-0.0017*((([.F37]+273.15)/100))^2))","v":5.45709556867941},null,4.0834]
["Latitude","03º59,87 N",null,null,116,20.9577,22.6,{"$f":"of:=[.G38]+273","v":295.6},35.8698,1.33,{"$f":"of:=([.J$135]*[.J38]*112/[.J$136])*([.H38]/273)","v":3.04898840219766},{"$f":"of:=[.K38]/[.M38]","v":{"v":0.603779055217878,"d":"60%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F38]+273.15))+143.3483*LN(([.F38]+273.15)/100)-21.8492*(([.F38]+273.15)/100)+[.I38]*(-0.033096+0.014259*(([.F38]+273.15)/100)-0.0017*((([.F38]+273.15)/100))^2))","v":5.04984128854455},null,3.0899]
["Longitude","037º59,73 W",null,null,399,10.0437,16.1,{"$f":"of:=[.G39]+273","v":289.1},34.9051,1.37,{"$f":"of:=([.J$135]*[.J39]*112/[.J$136])*([.H39]/273)","v":3.07162618019988},{"$f":"of:=[.K39]/[.M39]","v":{"v":0.486304539585351,"d":"49%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F39]+273.15))+143.3483*LN(([.F39]+273.15)/100)-21.8492*(([.F39]+273.15)/100)+[.I39]*(-0.033096+0.014259*(([.F39]+273.15)/100)-0.0017*((([.F39]+273.15)/100))^2))","v":6.31626055314825},null,3.1249]
[{"$nulls":4},900,5.5635,14,{"$f":"of:=[.G40]+273","v":287},34.559,1.3,{"$f":"of:=([.J$135]*[.J40]*112/[.J$136])*([.H40]/273)","v":2.89350975424},{"$f":"of:=[.K40]/[.M40]","v":{"v":0.41251510407042,"d":"41%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F40]+273.15))+143.3483*LN(([.F40]+273.15)/100)-21.8492*(([.F40]+273.15)/100)+[.I40]*(-0.033096+0.014259*(([.F40]+273.15)/100)-0.0017*((([.F40]+273.15)/100))^2))","v":7.01431226563294},null,2.9003]
[{"$nulls":4},3512,2.4395,12.7,{"$f":"of:=[.G41]+273","v":285.7},34.9027,2.6,{"$f":"of:=([.J$135]*[.J41]*112/[.J$136])*([.H41]/273)","v":5.760806528128},{"$f":"of:=[.K41]/[.M41]","v":{"v":0.762309084862208,"d":"76%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F41]+273.15))+143.3483*LN(([.F41]+273.15)/100)-21.8492*(([.F41]+273.15)/100)+[.I41]*(-0.033096+0.014259*(([.F41]+273.15)/100)-0.0017*((([.F41]+273.15)/100))^2))","v":7.55704824004465},null,5.705]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[15,"2011-08-29","PT00H24M00S",4645,4,28.6322,26.2,{"$f":"of:=[.G44]+273","v":299.2},35.7642,1.92,{"$f":"of:=([.J$135]*[.J44]*112/[.J$136])*([.H44]/273)","v":4.45515193637022},{"$f":"of:=[.K44]/[.M44]","v":{"v":1.00550629299989,"d":"101%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F44]+273.15))+143.3483*LN(([.F44]+273.15)/100)-21.8492*(([.F44]+273.15)/100)+[.I44]*(-0.033096+0.014259*(([.F44]+273.15)/100)-0.0017*((([.F44]+273.15)/100))^2))","v":4.43075490167092},null,4.2879]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},98,19.4434,21.4,{"$f":"of:=[.G45]+273","v":294.4},36.5786,1.25,{"$f":"of:=([.J$135]*[.J45]*112/[.J$136])*([.H45]/273)","v":2.85395762412308},{"$f":"of:=[.K45]/[.M45]","v":{"v":0.55195529361467,"d":"55%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F45]+273.15))+143.3483*LN(([.F45]+273.15)/100)-21.8492*(([.F45]+273.15)/100)+[.I45]*(-0.033096+0.014259*(([.F45]+273.15)/100)-0.0017*((([.F45]+273.15)/100))^2))","v":5.17063185576671},null,2.888]
["Latitude","11º30,31 N",null,null,178,12.1489,17.4,{"$f":"of:=[.G46]+273","v":290.4},35.275,1.02,{"$f":"of:=([.J$135]*[.J46]*112/[.J$136])*([.H46]/273)","v":2.29718771719089},{"$f":"of:=[.K46]/[.M46]","v":{"v":0.381453837794379,"d":"38%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F46]+273.15))+143.3483*LN(([.F46]+273.15)/100)-21.8492*(([.F46]+273.15)/100)+[.I46]*(-0.033096+0.014259*(([.F46]+273.15)/100)-0.0017*((([.F46]+273.15)/100))^2))","v":6.02219060233753},null,2.1909]
["Longitude","037º5720 W",null,null,698,6.8656,14.6,{"$f":"of:=[.G47]+273","v":287.6},34.7975,0.98,{"$f":"of:=([.J$135]*[.J47]*112/[.J$136])*([.H47]/273)","v":2.18582132729305},{"$f":"of:=[.K47]/[.M47]","v":{"v":0.321840117243715,"d":"32%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F47]+273.15))+143.3483*LN(([.F47]+273.15)/100)-21.8492*(([.F47]+273.15)/100)+[.I47]*(-0.033096+0.014259*(([.F47]+273.15)/100)-0.0017*((([.F47]+273.15)/100))^2))","v":6.79163724526556},null,2.0995]
[{"$nulls":4},1500,4.3344,15.6,{"$f":"of:=[.G48]+273","v":288.6},34.9772,2.21,{"$f":"of:=([.J$135]*[.J48]*112/[.J$136])*([.H48]/273)","v":4.9463893924224},{"$f":"of:=[.K48]/[.M48]","v":{"v":0.68645563844682,"d":"69%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F48]+273.15))+143.3483*LN(([.F48]+273.15)/100)-21.8492*(([.F48]+273.15)/100)+[.I48]*(-0.033096+0.014259*(([.F48]+273.15)/100)-0.0017*((([.F48]+273.15)/100))^2))","v":7.20569417073205},null,4.9432]
[{"$nulls":4},4003,2.3963,14,{"$f":"of:=[.G49]+273","v":287},34.8895,2.36,{"$f":"of:=([.J$135]*[.J49]*112/[.J$136])*([.H49]/273)","v":5.25283309231262},{"$f":"of:=[.K49]/[.M49]","v":{"v":0.694268919951057,"d":"69%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F49]+273.15))+143.3483*LN(([.F49]+273.15)/100)-21.8492*(([.F49]+273.15)/100)+[.I49]*(-0.033096+0.014259*(([.F49]+273.15)/100)-0.0017*((([.F49]+273.15)/100))^2))","v":7.56599199728387},null,5.4623]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[17,"2011-08-29","PT16H11M00S",5368,2,28.211,30,{"$f":"of:=[.G52]+273","v":303},36.3471,1.87,{"$f":"of:=([.J$135]*[.J52]*112/[.J$136])*([.H52]/273)","v":4.39424165599015},{"$f":"of:=[.K52]/[.M52]","v":{"v":0.988210876885299,"d":"99%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F52]+273.15))+143.3483*LN(([.F52]+273.15)/100)-21.8492*(([.F52]+273.15)/100)+[.I52]*(-0.033096+0.014259*(([.F52]+273.15)/100)-0.0017*((([.F52]+273.15)/100))^2))","v":4.44666392444514},null,4.33]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},1,28.21,30,{"$f":"of:=[.G53]+273","v":303},36.3522,1.86,{"$f":"of:=([.J$135]*[.J53]*112/[.J$136])*([.H53]/273)","v":4.37074303750892},{"$f":"of:=[.K53]/[.M53]","v":{"v":0.982938261242481,"d":"98%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F53]+273.15))+143.3483*LN(([.F53]+273.15)/100)-21.8492*(([.F53]+273.15)/100)+[.I53]*(-0.033096+0.014259*(([.F53]+273.15)/100)-0.0017*((([.F53]+273.15)/100))^2))","v":4.44660993456913},null,4.3351]
["Latitude","13º14,43 N",null,null,17,27.88,28.8,{"$f":"of:=[.G54]+273","v":301.8},36.4313,1.86,{"$f":"of:=([.J$135]*[.J54]*112/[.J$136])*([.H54]/273)","v":4.35343316409305},{"$f":"of:=[.K54]/[.M54]","v":{"v":0.974231413061641,"d":"97%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F54]+273.15))+143.3483*LN(([.F54]+273.15)/100)-21.8492*(([.F54]+273.15)/100)+[.I54]*(-0.033096+0.014259*(([.F54]+273.15)/100)-0.0017*((([.F54]+273.15)/100))^2))","v":4.46858221334893},null,4.3624]
["Longitude","038º00,22 W",null,null,142,14.92,21.2,{"$f":"of:=[.G55]+273","v":294.2},35.8174,1.03,{"$f":"of:=([.J$135]*[.J55]*112/[.J$136])*([.H55]/273)","v":2.35006348643348},{"$f":"of:=[.K55]/[.M55]","v":{"v":0.414554347211075,"d":"41%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F55]+273.15))+143.3483*LN(([.F55]+273.15)/100)-21.8492*(([.F55]+273.15)/100)+[.I55]*(-0.033096+0.014259*(([.F55]+273.15)/100)-0.0017*((([.F55]+273.15)/100))^2))","v":5.66889118940276},null,2.1911]
[{"$nulls":4},597,8.2,16.6,{"$f":"of:=[.G56]+273","v":289.6},34.9226,0.82,{"$f":"of:=([.J$135]*[.J56]*112/[.J$136])*([.H56]/273)","v":1.84167126335803},{"$f":"of:=[.K56]/[.M56]","v":{"v":0.27984505756958,"d":"28%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F56]+273.15))+143.3483*LN(([.F56]+273.15)/100)-21.8492*(([.F56]+273.15)/100)+[.I56]*(-0.033096+0.014259*(([.F56]+273.15)/100)-0.0017*((([.F56]+273.15)/100))^2))","v":6.58103909124827},null,1.8393]
[{"$nulls":4},1498,4.42,14.4,{"$f":"of:=[.G57]+273","v":287.4},34.9922,2.07,{"$f":"of:=([.J$135]*[.J57]*112/[.J$136])*([.H57]/273)","v":4.61377924409797},{"$f":"of:=[.K57]/[.M57]","v":{"v":0.641698442552378,"d":"64%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F57]+273.15))+143.3483*LN(([.F57]+273.15)/100)-21.8492*(([.F57]+273.15)/100)+[.I57]*(-0.033096+0.014259*(([.F57]+273.15)/100)-0.0017*((([.F57]+273.15)/100))^2))","v":7.18994926300038},null,4.7925]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[19,"2011-08-30","PT05H11M00S",5764,1,27.4486,25.9,{"$f":"of:=[.G60]+273","v":298.9},36.6566,1.98,{"$f":"of:=([.J$135]*[.J60]*112/[.J$136])*([.H60]/273)","v":4.58976877452111},{"$f":"of:=[.K60]/[.M60]","v":{"v":1.02118224668902,"d":"102%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F60]+273.15))+143.3483*LN(([.F60]+273.15)/100)-21.8492*(([.F60]+273.15)/100)+[.I60]*(-0.033096+0.014259*(([.F60]+273.15)/100)-0.0017*((([.F60]+273.15)/100))^2))","v":4.49456381503155},null,4.352]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},32,25.99,25.7,{"$f":"of:=[.G61]+273","v":298.7},36.7947,1.96,{"$f":"of:=([.J$135]*[.J61]*112/[.J$136])*([.H61]/273)","v":4.54036738847311},{"$f":"of:=[.K61]/[.M61]","v":{"v":0.98682398412286,"d":"99%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F61]+273.15))+143.3483*LN(([.F61]+273.15)/100)-21.8492*(([.F61]+273.15)/100)+[.I61]*(-0.033096+0.014259*(([.F61]+273.15)/100)-0.0017*((([.F61]+273.15)/100))^2))","v":4.60099010717582},null,4.6497]
["Latitude","14º58,19 N",null,null,55,24.5089,24.7,{"$f":"of:=[.G62]+273","v":297.7},36.9947,2.05,{"$f":"of:=([.J$135]*[.J62]*112/[.J$136])*([.H62]/273)","v":4.732955240864},{"$f":"of:=[.K62]/[.M62]","v":{"v":1.0043362347715,"d":"100%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F62]+273.15))+143.3483*LN(([.F62]+273.15)/100)-21.8492*(([.F62]+273.15)/100)+[.I62]*(-0.033096+0.014259*(([.F62]+273.15)/100)-0.0017*((([.F62]+273.15)/100))^2))","v":4.7125206449818},null,4.7343]
["Longitude","038º00,34 W",null,null,203,14.2858,19.2,{"$f":"of:=[.G63]+273","v":292.2},35.7775,1.12,{"$f":"of:=([.J$135]*[.J63]*112/[.J$136])*([.H63]/273)","v":2.53803692364406},{"$f":"of:=[.K63]/[.M63]","v":{"v":0.441903170132548,"d":"44%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F63]+273.15))+143.3483*LN(([.F63]+273.15)/100)-21.8492*(([.F63]+273.15)/100)+[.I63]*(-0.033096+0.014259*(([.F63]+273.15)/100)-0.0017*((([.F63]+273.15)/100))^2))","v":5.74342320939399},null,2.2023]
[{"$nulls":4},1502,4.534,14,{"$f":"of:=[.G64]+273","v":287},34.992,2.02,{"$f":"of:=([.J$135]*[.J64]*112/[.J$136])*([.H64]/273)","v":4.49606900274215},{"$f":"of:=[.K64]/[.M64]","v":{"v":0.627064256728998,"d":"63%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F64]+273.15))+143.3483*LN(([.F64]+273.15)/100)-21.8492*(([.F64]+273.15)/100)+[.I64]*(-0.033096+0.014259*(([.F64]+273.15)/100)-0.0017*((([.F64]+273.15)/100))^2))","v":7.17002915489926},null,4.6521]
[{"$nulls":4},4080,2.3828,14.2,{"$f":"of:=[.G65]+273","v":287.2},34.8886,2.44,{"$f":"of:=([.J$135]*[.J65]*112/[.J$136])*([.H65]/273)","v":5.43467982701489},{"$f":"of:=[.K65]/[.M65]","v":{"v":0.718053622832221,"d":"72%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F65]+273.15))+143.3483*LN(([.F65]+273.15)/100)-21.8492*(([.F65]+273.15)/100)+[.I65]*(-0.033096+0.014259*(([.F65]+273.15)/100)-0.0017*((([.F65]+273.15)/100))^2))","v":7.56862670726299},null,5.4402]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[24,"2011-09-14","PT21H52M00S",4520,1,27.41,26,{"$f":"of:=[.G68]+273","v":299},35.5338,1.94,{"$f":"of:=([.J$135]*[.J68]*112/[.J$136])*([.H68]/273)","v":4.498550705024},{"$f":"of:=[.K68]/[.M68]","v":{"v":0.993987920048255,"d":"99%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F68]+273.15))+143.3483*LN(([.F68]+273.15)/100)-21.8492*(([.F68]+273.15)/100)+[.I68]*(-0.033096+0.014259*(([.F68]+273.15)/100)-0.0017*((([.F68]+273.15)/100))^2))","v":4.52575993559923},null,4.3444]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},106,24.2355,25.1,{"$f":"of:=[.G69]+273","v":298.1},36.6942,1.82,{"$f":"of:=([.J$135]*[.J69]*112/[.J$136])*([.H69]/273)","v":4.2075866231168},{"$f":"of:=[.K69]/[.M69]","v":{"v":0.887145176469739,"d":"89%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F69]+273.15))+143.3483*LN(([.F69]+273.15)/100)-21.8492*(([.F69]+273.15)/100)+[.I69]*(-0.033096+0.014259*(([.F69]+273.15)/100)-0.0017*((([.F69]+273.15)/100))^2))","v":4.74283886641898},null,4.1326]
["Latitude","00º00,81 S",null,null,501,7.18,15.9,{"$f":"of:=[.G70]+273","v":288.9},34.6203,1.38,{"$f":"of:=([.J$135]*[.J70]*112/[.J$136])*([.H70]/273)","v":3.09190634103434},{"$f":"of:=[.K70]/[.M70]","v":{"v":0.458061196551878,"d":"46%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F70]+273.15))+143.3483*LN(([.F70]+273.15)/100)-21.8492*(([.F70]+273.15)/100)+[.I70]*(-0.033096+0.014259*(([.F70]+273.15)/100)-0.0017*((([.F70]+273.15)/100))^2))","v":6.74998529521625},null,2.9938]
["Longitude","035°01,93W",null,null,762,4.83,14.3,{"$f":"of:=[.G71]+273","v":287.3},34.4962,1.65,{"$f":"of:=([.J$135]*[.J71]*112/[.J$136])*([.H71]/273)","v":3.676370494368},{"$f":"of:=[.K71]/[.M71]","v":{"v":0.514749740489186,"d":"51%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F71]+273.15))+143.3483*LN(([.F71]+273.15)/100)-21.8492*(([.F71]+273.15)/100)+[.I71]*(-0.033096+0.014259*(([.F71]+273.15)/100)-0.0017*((([.F71]+273.15)/100))^2))","v":7.1420540996761},null,3.5222]
[{"$nulls":4},1507,4.29,14,{"$f":"of:=[.G72]+273","v":287},34.9764,2.36,{"$f":"of:=([.J$135]*[.J72]*112/[.J$136])*([.H72]/273)","v":5.25283309231262},{"$f":"of:=[.K72]/[.M72]","v":{"v":0.728189630044794,"d":"73%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F72]+273.15))+143.3483*LN(([.F72]+273.15)/100)-21.8492*(([.F72]+273.15)/100)+[.I72]*(-0.033096+0.014259*(([.F72]+273.15)/100)-0.0017*((([.F72]+273.15)/100))^2))","v":7.21355108007991},null,5.1901]
[{"$nulls":4},4010,2.04,14,{"$f":"of:=[.G73]+273","v":287},34.8585,2.5,{"$f":"of:=([.J$135]*[.J73]*112/[.J$136])*([.H73]/273)","v":5.56444183507692},{"$f":"of:=[.K73]/[.M73]","v":{"v":0.728669596940358,"d":"73%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F73]+273.15))+143.3483*LN(([.F73]+273.15)/100)-21.8492*(([.F73]+273.15)/100)+[.I73]*(-0.033096+0.014259*(([.F73]+273.15)/100)-0.0017*((([.F73]+273.15)/100))^2))","v":7.63644024457957},null,5.5508]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[25,"2011-09-18","PT09H01M00S",5365,1,26.47,26,{"$f":"of:=[.G76]+273","v":299},36.1935,1.89,{"$f":"of:=([.J$135]*[.J76]*112/[.J$136])*([.H76]/273)","v":4.382608676544},{"$f":"of:=[.K76]/[.M76]","v":{"v":0.956952624360357,"d":"96%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F76]+273.15))+143.3483*LN(([.F76]+273.15)/100)-21.8492*(([.F76]+273.15)/100)+[.I76]*(-0.033096+0.014259*(([.F76]+273.15)/100)-0.0017*((([.F76]+273.15)/100))^2))","v":4.57975511533124},null,4.3727]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},93,25.35,25.1,{"$f":"of:=[.G77]+273","v":298.1},36.2656,1.89,{"$f":"of:=([.J$135]*[.J77]*112/[.J$136])*([.H77]/273)","v":4.36941687785206},{"$f":"of:=[.K77]/[.M77]","v":{"v":0.936672146400205,"d":"94%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F77]+273.15))+143.3483*LN(([.F77]+273.15)/100)-21.8492*(([.F77]+273.15)/100)+[.I77]*(-0.033096+0.014259*(([.F77]+273.15)/100)-0.0017*((([.F77]+273.15)/100))^2))","v":4.66483058628838},null,4.2243]
["Latitude","07°49,13 S",null,null,706,5.32,16.8,{"$f":"of:=[.G78]+273","v":289.8},34.477,1.44,{"$f":"of:=([.J$135]*[.J78]*112/[.J$136])*([.H78]/273)","v":3.23638794575557},{"$f":"of:=[.K78]/[.M78]","v":{"v":0.458469970206153,"d":"46%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F78]+273.15))+143.3483*LN(([.F78]+273.15)/100)-21.8492*(([.F78]+273.15)/100)+[.I78]*(-0.033096+0.014259*(([.F78]+273.15)/100)-0.0017*((([.F78]+273.15)/100))^2))","v":7.05910562539202},null,3.2687]
["Longitude","030°32,07 W",null,null,1111,4.11,14.8,{"$f":"of:=[.G79]+273","v":287.8},34.6561,1.75,{"$f":"of:=([.J$135]*[.J79]*112/[.J$136])*([.H79]/273)","v":3.90596673203692},{"$f":"of:=[.K79]/[.M79]","v":{"v":0.537952215947332,"d":"54%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F79]+273.15))+143.3483*LN(([.F79]+273.15)/100)-21.8492*(([.F79]+273.15)/100)+[.I79]*(-0.033096+0.014259*(([.F79]+273.15)/100)-0.0017*((([.F79]+273.15)/100))^2))","v":7.26080610181803},null,3.837]
[{"$nulls":4},2022,3.37,14.2,{"$f":"of:=[.G80]+273","v":287.2},34.9443,2.42,{"$f":"of:=([.J$135]*[.J80]*112/[.J$136])*([.H80]/273)","v":5.39013327105575},{"$f":"of:=[.K80]/[.M80]","v":{"v":0.730311747104554,"d":"73%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F80]+273.15))+143.3483*LN(([.F80]+273.15)/100)-21.8492*(([.F80]+273.15)/100)+[.I80]*(-0.033096+0.014259*(([.F80]+273.15)/100)-0.0017*((([.F80]+273.15)/100))^2))","v":7.38059231886364},null,5.4172]
[{"$nulls":4},4060,1.8,13.6,{"$f":"of:=[.G81]+273","v":286.6},34.0291,2.38,{"$f":"of:=([.J$135]*[.J81]*112/[.J$136])*([.H81]/273)","v":5.28996556270474},{"$f":"of:=[.K81]/[.M81]","v":{"v":0.684648425320404,"d":"68%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F81]+273.15))+143.3483*LN(([.F81]+273.15)/100)-21.8492*(([.F81]+273.15)/100)+[.I81]*(-0.033096+0.014259*(([.F81]+273.15)/100)-0.0017*((([.F81]+273.15)/100))^2))","v":7.72654309433213},null,5.343]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[27,"2011-09-29","PT10H32M00S",5105,1,25.95,25.9,{"$f":"of:=[.G84]+273","v":298.9},36.4604,1.9,{"$f":"of:=([.J$135]*[.J84]*112/[.J$136])*([.H84]/273)","v":4.40432357151015},{"$f":"of:=[.K84]/[.M84]","v":{"v":0.954807143345101,"d":"95%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F84]+273.15))+143.3483*LN(([.F84]+273.15)/100)-21.8492*(([.F84]+273.15)/100)+[.I84]*(-0.033096+0.014259*(([.F84]+273.15)/100)-0.0017*((([.F84]+273.15)/100))^2))","v":4.61278866858904},null,4.4301]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},81,25.74,25.7,{"$f":"of:=[.G85]+273","v":298.7},36.623,1.92,{"$f":"of:=([.J$135]*[.J85]*112/[.J$136])*([.H85]/273)","v":4.44770682952468},{"$f":"of:=[.K85]/[.M85]","v":{"v":0.961700786876077,"d":"96%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F85]+273.15))+143.3483*LN(([.F85]+273.15)/100)-21.8492*(([.F85]+273.15)/100)+[.I85]*(-0.033096+0.014259*(([.F85]+273.15)/100)-0.0017*((([.F85]+273.15)/100))^2))","v":4.6248343457972},null,4.4406]
["Latitude","10°42,80 S",null,null,132,21.73,22.7,{"$f":"of:=[.G86]+273","v":295.7},36.6165,1.81,{"$f":"of:=([.J$135]*[.J86]*112/[.J$136])*([.H86]/273)","v":4.15077890682142},{"$f":"of:=[.K86]/[.M86]","v":{"v":0.83713617104065,"d":"84%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F86]+273.15))+143.3483*LN(([.F86]+273.15)/100)-21.8492*(([.F86]+273.15)/100)+[.I86]*(-0.033096+0.014259*(([.F86]+273.15)/100)-0.0017*((([.F86]+273.15)/100))^2))","v":4.95830791979942},null,4.1153]
["Longitude","031°31,4 W",null,null,308,10.1,15.1,{"$f":"of:=[.G87]+273","v":288.1},34.9312,1.03,{"$f":"of:=([.J$135]*[.J87]*112/[.J$136])*([.H87]/273)","v":2.30133681319335},{"$f":"of:=[.K87]/[.M87]","v":{"v":0.364862108264956,"d":"36%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F87]+273.15))+143.3483*LN(([.F87]+273.15)/100)-21.8492*(([.F87]+273.15)/100)+[.I87]*(-0.033096+0.014259*(([.F87]+273.15)/100)-0.0017*((([.F87]+273.15)/100))^2))","v":6.30741521540013},null,2.205]
[{"$nulls":4},703,5.34,15.4,{"$f":"of:=[.G88]+273","v":288.4},34.476,1.46,{"$f":"of:=([.J$135]*[.J88]*112/[.J$136])*([.H88]/273)","v":3.26548590501022},{"$f":"of:=[.K88]/[.M88]","v":{"v":0.462810921964643,"d":"46%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F88]+273.15))+143.3483*LN(([.F88]+273.15)/100)-21.8492*(([.F88]+273.15)/100)+[.I88]*(-0.033096+0.014259*(([.F88]+273.15)/100)-0.0017*((([.F88]+273.15)/100))^2))","v":7.05576672898762},null,3.2631]
[{"$nulls":4},1488,4.23,14.2,{"$f":"of:=[.G89]+273","v":287.2},34.9193,2.15,{"$f":"of:=([.J$135]*[.J89]*112/[.J$136])*([.H89]/273)","v":4.78875476560738},{"$f":"of:=[.K89]/[.M89]","v":{"v":0.662631565832609,"d":"66%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F89]+273.15))+143.3483*LN(([.F89]+273.15)/100)-21.8492*(([.F89]+273.15)/100)+[.I89]*(-0.033096+0.014259*(([.F89]+273.15)/100)-0.0017*((([.F89]+273.15)/100))^2))","v":7.22687389573755},null,4.7985]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[28,"2011-09-29","PT22H10M00S",2587,6,25.8,25.6,{"$f":"of:=[.G92]+273","v":298.6},36.5752,1.92,{"$f":"of:=([.J$135]*[.J92]*112/[.J$136])*([.H92]/273)","v":4.44621780815557},{"$f":"of:=[.K92]/[.M92]","v":{"v":0.962088947826528,"d":"96%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F92]+273.15))+143.3483*LN(([.F92]+273.15)/100)-21.8492*(([.F92]+273.15)/100)+[.I92]*(-0.033096+0.014259*(([.F92]+273.15)/100)-0.0017*((([.F92]+273.15)/100))^2))","v":4.62142073058848},null,4.4532]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},134,22.78,23,{"$f":"of:=[.G93]+273","v":296},36.8273,1.88,{"$f":"of:=([.J$135]*[.J93]*112/[.J$136])*([.H93]/273)","v":4.31568026813046},{"$f":"of:=[.K93]/[.M93]","v":{"v":0.88785563879352,"d":"89%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F93]+273.15))+143.3483*LN(([.F93]+273.15)/100)-21.8492*(([.F93]+273.15)/100)+[.I93]*(-0.033096+0.014259*(([.F93]+273.15)/100)-0.0017*((([.F93]+273.15)/100))^2))","v":4.86079051544337},null,4.3473]
["Latitude","12°08,30 S",null,null,306,11.53,16.2,{"$f":"of:=[.G94]+273","v":289.2},35.0771,1.13,{"$f":"of:=([.J$135]*[.J94]*112/[.J$136])*([.H94]/273)","v":2.53440743405686},{"$f":"of:=[.K94]/[.M94]","v":{"v":0.414831754195133,"d":"41%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F94]+273.15))+143.3483*LN(([.F94]+273.15)/100)-21.8492*(([.F94]+273.15)/100)+[.I94]*(-0.033096+0.014259*(([.F94]+273.15)/100)-0.0017*((([.F94]+273.15)/100))^2))","v":6.10948272022758},null,2.5138]
["Longitude","032°03,37 W",null,null,504,6.62,15.8,{"$f":"of:=[.G95]+273","v":288.8},34.5373,1.48,{"$f":"of:=([.J$135]*[.J95]*112/[.J$136])*([.H95]/273)","v":3.31480973786191},{"$f":"of:=[.K95]/[.M95]","v":{"v":0.484460964458959,"d":"48%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F95]+273.15))+143.3483*LN(([.F95]+273.15)/100)-21.8492*(([.F95]+273.15)/100)+[.I95]*(-0.033096+0.014259*(([.F95]+273.15)/100)-0.0017*((([.F95]+273.15)/100))^2))","v":6.8422638376321},null,3.324]
[{"$nulls":4},709,5.07,15,{"$f":"of:=[.G96]+273","v":288},34.4534,1.55,{"$f":"of:=([.J$135]*[.J96]*112/[.J$136])*([.H96]/273)","v":3.46197468317538},{"$f":"of:=[.K96]/[.M96]","v":{"v":0.487411864479719,"d":"49%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F96]+273.15))+143.3483*LN(([.F96]+273.15)/100)-21.8492*(([.F96]+273.15)/100)+[.I96]*(-0.033096+0.014259*(([.F96]+273.15)/100)-0.0017*((([.F96]+273.15)/100))^2))","v":7.10277064525464},null,3.4634]
[{"$nulls":4},1517,4.09,14.7,{"$f":"of:=[.G97]+273","v":287.7},34.9037,2.18,{"$f":"of:=([.J$135]*[.J97]*112/[.J$136])*([.H97]/273)","v":4.86402789794363},{"$f":"of:=[.K97]/[.M97]","v":{"v":0.670673427432772,"d":"67%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F97]+273.15))+143.3483*LN(([.F97]+273.15)/100)-21.8492*(([.F97]+273.15)/100)+[.I97]*(-0.033096+0.014259*(([.F97]+273.15)/100)-0.0017*((([.F97]+273.15)/100))^2))","v":7.25245357723853},null,4.7825]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[29,"2011-09-30","PT09H15M00S",4670,6,25.41,25.5,{"$f":"of:=[.G100]+273","v":298.5},36.927,1.95,{"$f":"of:=([.J$135]*[.J100]*112/[.J$136])*([.H100]/273)","v":4.51417767408},{"$f":"of:=[.K100]/[.M100]","v":{"v":0.972332609574864,"d":"97%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F100]+273.15))+143.3483*LN(([.F100]+273.15)/100)-21.8492*(([.F100]+273.15)/100)+[.I100]*(-0.033096+0.014259*(([.F100]+273.15)/100)-0.0017*((([.F100]+273.15)/100))^2))","v":4.6426270492499},null,4.4874]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},120,25.24,25.2,{"$f":"of:=[.G101]+273","v":298.2},37.0082,1.93,{"$f":"of:=([.J$135]*[.J101]*112/[.J$136])*([.H101]/273)","v":4.46338808581809},{"$f":"of:=[.K101]/[.M101]","v":{"v":0.959074405413272,"d":"96%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F101]+273.15))+143.3483*LN(([.F101]+273.15)/100)-21.8492*(([.F101]+273.15)/100)+[.I101]*(-0.033096+0.014259*(([.F101]+273.15)/100)-0.0017*((([.F101]+273.15)/100))^2))","v":4.65384964985567},null,4.4698]
["Latitude","13°33,10 S",null,null,710,4.79,16,{"$f":"of:=[.G102]+273","v":289},34.3764,1.67,{"$f":"of:=([.J$135]*[.J102]*112/[.J$136])*([.H102]/273)","v":3.74294991339815},{"$f":"of:=[.K102]/[.M102]","v":{"v":0.52315081926894,"d":"52%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F102]+273.15))+143.3483*LN(([.F102]+273.15)/100)-21.8492*(([.F102]+273.15)/100)+[.I102]*(-0.033096+0.014259*(([.F102]+273.15)/100)-0.0017*((([.F102]+273.15)/100))^2))","v":7.15462879065852},null,3.8818]
["Longitude","032°31,80 W",null,null,1107,3.92,14.5,{"$f":"of:=[.G103]+273","v":287.5},34.6534,1.83,{"$f":"of:=([.J$135]*[.J103]*112/[.J$136])*([.H103]/273)","v":4.08026754073846},{"$f":"of:=[.K103]/[.M103]","v":{"v":0.559329656960069,"d":"56%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F103]+273.15))+143.3483*LN(([.F103]+273.15)/100)-21.8492*(([.F103]+273.15)/100)+[.I103]*(-0.033096+0.014259*(([.F103]+273.15)/100)-0.0017*((([.F103]+273.15)/100))^2))","v":7.29492436162697},null,4.0764]
[{"$nulls":4},2019,3.29,13.9,{"$f":"of:=[.G104]+273","v":286.9},34.9471,2.46,{"$f":"of:=([.J$135]*[.J104]*112/[.J$136])*([.H104]/273)","v":5.47350295708652},{"$f":"of:=[.K104]/[.M104]","v":{"v":0.740146275895777,"d":"74%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F104]+273.15))+143.3483*LN(([.F104]+273.15)/100)-21.8492*(([.F104]+273.15)/100)+[.I104]*(-0.033096+0.014259*(([.F104]+273.15)/100)-0.0017*((([.F104]+273.15)/100))^2))","v":7.3951638146961},null,5.5305]
[{"$nulls":4},4062,1.71,13.2,{"$f":"of:=[.G105]+273","v":286.2},34.821,2.38,{"$f":"of:=([.J$135]*[.J105]*112/[.J$136])*([.H105]/273)","v":5.28258249841625},{"$f":"of:=[.K105]/[.M105]","v":{"v":0.685767735035112,"d":"69%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F105]+273.15))+143.3483*LN(([.F105]+273.15)/100)-21.8492*(([.F105]+273.15)/100)+[.I105]*(-0.033096+0.014259*(([.F105]+273.15)/100)-0.0017*((([.F105]+273.15)/100))^2))","v":7.70316570543491},null,5.2941]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[30,"2011-10-01","PT07H30M00S",4535,1,25.0499,25.1,{"$f":"of:=[.G108]+273","v":298.1},37.1042,1.98,{"$f":"of:=([.J$135]*[.J108]*112/[.J$136])*([.H108]/273)","v":4.57748434822597},{"$f":"of:=[.K108]/[.M108]","v":{"v":0.980960713680344,"d":"98%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F108]+273.15))+143.3483*LN(([.F108]+273.15)/100)-21.8492*(([.F108]+273.15)/100)+[.I108]*(-0.033096+0.014259*(([.F108]+273.15)/100)-0.0017*((([.F108]+273.15)/100))^2))","v":4.66632790119828},null,4.4745]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},115,24.1443,24.9,{"$f":"of:=[.G109]+273","v":297.9},37.2177,1.95,{"$f":"of:=([.J$135]*[.J109]*112/[.J$136])*([.H109]/273)","v":4.505103950112},{"$f":"of:=[.K109]/[.M109]","v":{"v":0.951233269717287,"d":"95%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F109]+273.15))+143.3483*LN(([.F109]+273.15)/100)-21.8492*(([.F109]+273.15)/100)+[.I109]*(-0.033096+0.014259*(([.F109]+273.15)/100)-0.0017*((([.F109]+273.15)/100))^2))","v":4.73606642401285},null,4.5146]
["Latitude","14°53,90 S",null,null,253,14.6943,18.2,{"$f":"of:=[.G110]+273","v":291.2},35.4133,1.77,{"$f":"of:=([.J$135]*[.J110]*112/[.J$136])*([.H110]/273)","v":3.9972778653696},{"$f":"of:=[.K110]/[.M110]","v":{"v":0.700189008522,"d":"70%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F110]+273.15))+143.3483*LN(([.F110]+273.15)/100)-21.8492*(([.F110]+273.15)/100)+[.I110]*(-0.033096+0.014259*(([.F110]+273.15)/100)-0.0017*((([.F110]+273.15)/100))^2))","v":5.70885549004445},null,3.9912]
["Longitude","033°07,20 W",null,null,533,6.7926,16.4,{"$f":"of:=[.G111]+273","v":289.4},34.5585,1.41,{"$f":"of:=([.J$135]*[.J111]*112/[.J$136])*([.H111]/273)","v":3.16458919661391},{"$f":"of:=[.K111]/[.M111]","v":{"v":0.46443954894104,"d":"46%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F111]+273.15))+143.3483*LN(([.F111]+273.15)/100)-21.8492*(([.F111]+273.15)/100)+[.I111]*(-0.033096+0.014259*(([.F111]+273.15)/100)-0.0017*((([.F111]+273.15)/100))^2))","v":6.81378061758398},null,3.166]
[{"$nulls":4},712,4.5853,16,{"$f":"of:=[.G112]+273","v":289},34.4036,1.82,{"$f":"of:=([.J$135]*[.J112]*112/[.J$136])*([.H112]/273)","v":4.079143019392},{"$f":"of:=[.K112]/[.M112]","v":{"v":0.567415614691754,"d":"57%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F112]+273.15))+143.3483*LN(([.F112]+273.15)/100)-21.8492*(([.F112]+273.15)/100)+[.I112]*(-0.033096+0.014259*(([.F112]+273.15)/100)-0.0017*((([.F112]+273.15)/100))^2))","v":7.1889861924367},null,4.0632]
[{"$nulls":4},1502,3.9239,14.8,{"$f":"of:=[.G113]+273","v":287.8},34.9044,2.2,{"$f":"of:=([.J$135]*[.J113]*112/[.J$136])*([.H113]/273)","v":4.91035817741785},{"$f":"of:=[.K113]/[.M113]","v":{"v":0.674307171901374,"d":"67%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F113]+273.15))+143.3483*LN(([.F113]+273.15)/100)-21.8492*(([.F113]+273.15)/100)+[.I113]*(-0.033096+0.014259*(([.F113]+273.15)/100)-0.0017*((([.F113]+273.15)/100))^2))","v":7.28207912066527},null,4.9013]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[31,"2011-10-01","PT17H56M00S",4525,1,24.95,25,{"$f":"of:=[.G116]+273","v":298},37.2498,2.02,{"$f":"of:=([.J$135]*[.J116]*112/[.J$136])*([.H116]/273)","v":4.66839220493785},{"$f":"of:=[.K116]/[.M116]","v":{"v":0.999574131878664,"d":"100%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F116]+273.15))+143.3483*LN(([.F116]+273.15)/100)-21.8492*(([.F116]+273.15)/100)+[.I116]*(-0.033096+0.014259*(([.F116]+273.15)/100)-0.0017*((([.F116]+273.15)/100))^2))","v":4.67038117139323},null,4.5225]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},161,23.72,24.2,{"$f":"of:=[.G117]+273","v":297.2},37.2516,1.96,{"$f":"of:=([.J$135]*[.J117]*112/[.J$136])*([.H117]/273)","v":4.51756674875865},{"$f":"of:=[.K117]/[.M117]","v":{"v":0.9470852069425,"d":"95%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F117]+273.15))+143.3483*LN(([.F117]+273.15)/100)-21.8492*(([.F117]+273.15)/100)+[.I117]*(-0.033096+0.014259*(([.F117]+273.15)/100)-0.0017*((([.F117]+273.15)/100))^2))","v":4.7699686529186},null,4.4926]
["Latitude","16°14,21 S",null,null,199,19.93,20.8,{"$f":"of:=[.G118]+273","v":293.8},36.3533,1.88,{"$f":"of:=([.J$135]*[.J118]*112/[.J$136])*([.H118]/273)","v":4.2836042661376},{"$f":"of:=[.K118]/[.M118]","v":{"v":0.83483205844893,"d":"83%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F118]+273.15))+143.3483*LN(([.F118]+273.15)/100)-21.8492*(([.F118]+273.15)/100)+[.I118]*(-0.033096+0.014259*(([.F118]+273.15)/100)-0.0017*((([.F118]+273.15)/100))^2))","v":5.1310969946414},null,4.2721]
["Longitude","033°38,83 W",null,null,449,9.08,18.9,{"$f":"of:=[.G119]+273","v":291.9},34.7414,1.52,{"$f":"of:=([.J$135]*[.J119]*112/[.J$136])*([.H119]/273)","v":3.44094225633674},{"$f":"of:=[.K119]/[.M119]","v":{"v":0.532709576696487,"d":"53%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F119]+273.15))+143.3483*LN(([.F119]+273.15)/100)-21.8492*(([.F119]+273.15)/100)+[.I119]*(-0.033096+0.014259*(([.F119]+273.15)/100)-0.0017*((([.F119]+273.15)/100))^2))","v":6.45932118899606},null,3.4065]
[{"$nulls":4},702,4.67,14.5,{"$f":"of:=[.G120]+273","v":287.5},34.3923,1.88,{"$f":"of:=([.J$135]*[.J120]*112/[.J$136])*([.H120]/273)","v":4.19175026043077},{"$f":"of:=[.K120]/[.M120]","v":{"v":0.584237575813803,"d":"58%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F120]+273.15))+143.3483*LN(([.F120]+273.15)/100)-21.8492*(([.F120]+273.15)/100)+[.I120]*(-0.033096+0.014259*(([.F120]+273.15)/100)-0.0017*((([.F120]+273.15)/100))^2))","v":7.17473581631915},null,4.1837]
[{"$nulls":4},1557,4.01,13.2,{"$f":"of:=[.G121]+273","v":286.2},34.9352,2.26,{"$f":"of:=([.J$135]*[.J121]*112/[.J$136])*([.H121]/273)","v":5.01623380101711},{"$f":"of:=[.K121]/[.M121]","v":{"v":0.690447396277018,"d":"69%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F121]+273.15))+143.3483*LN(([.F121]+273.15)/100)-21.8492*(([.F121]+273.15)/100)+[.I121]*(-0.033096+0.014259*(([.F121]+273.15)/100)-0.0017*((([.F121]+273.15)/100))^2))","v":7.2651933051892},null,5.0429]
["ESTAÇÃO","DATA","HORA","PROFUND","PROFUND","TEMP (ºC)","TEMP (ºC)","TEMP (K)","SALINIDADE","VOL,gasto","WINKLER","PORCENTAGEM","SOLUBILIDADE",null,"CTD"]
[null,null,"HMG","LOCAL (m)","COLETA","CTD","BANCADA",null,"PSU","tiossulfato","OD (ml / L)",null,"DO OXIGENIO",null,"OD (ml / L)"]
[33,"2011-10-02","PT17H26M00S",4177,1,24.1,24.8,{"$f":"of:=[.G124]+273","v":297.8},37.315,1.97,{"$f":"of:=([.J$135]*[.J124]*112/[.J$136])*([.H124]/273)","v":4.54978234650486},{"$f":"of:=[.K124]/[.M124]","v":{"v":0.960469907878322,"d":"96%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F124]+273.15))+143.3483*LN(([.F124]+273.15)/100)-21.8492*(([.F124]+273.15)/100)+[.I124]*(-0.033096+0.014259*(([.F124]+273.15)/100)-0.0017*((([.F124]+273.15)/100))^2))","v":4.73703789070844},null,4.5539]
["POSIÇÃO GEOGRÁFICA",{"$nulls":3},138,23.77,24,{"$f":"of:=[.G125]+273","v":297},37.2781,1.94,{"$f":"of:=([.J$135]*[.J125]*112/[.J$136])*([.H125]/273)","v":4.46846006485662},{"$f":"of:=[.K125]/[.M125]","v":{"v":0.93774455673691,"d":"94%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F125]+273.15))+143.3483*LN(([.F125]+273.15)/100)-21.8492*(([.F125]+273.15)/100)+[.I125]*(-0.033096+0.014259*(([.F125]+273.15)/100)-0.0017*((([.F125]+273.15)/100))^2))","v":4.7651143723037},null,4.4991]
["Latitude","18°58,13 S",null,null,507,9.07,18.7,{"$f":"of:=[.G126]+273","v":291.7},34.7098,1.81,{"$f":"of:=([.J$135]*[.J126]*112/[.J$136])*([.H126]/273)","v":4.09463039269465},{"$f":"of:=[.K126]/[.M126]","v":{"v":0.633640551519807,"d":"63%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F126]+273.15))+143.3483*LN(([.F126]+273.15)/100)-21.8492*(([.F126]+273.15)/100)+[.I126]*(-0.033096+0.014259*(([.F126]+273.15)/100)-0.0017*((([.F126]+273.15)/100))^2))","v":6.46207125297386},null,4.0054]
["Longitude","034°47,72 W",null,null,1008,3.62,14.2,{"$f":"of:=[.G127]+273","v":287.2},34.4814,1.86,{"$f":"of:=([.J$135]*[.J127]*112/[.J$136])*([.H127]/273)","v":4.14282970419988},{"$f":"of:=[.K127]/[.M127]","v":{"v":0.563069393989177,"d":"56%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F127]+273.15))+143.3483*LN(([.F127]+273.15)/100)-21.8492*(([.F127]+273.15)/100)+[.I127]*(-0.033096+0.014259*(([.F127]+273.15)/100)-0.0017*((([.F127]+273.15)/100))^2))","v":7.35758282802262},null,4.1461]
[{"$nulls":4},2521,2.91,13.1,{"$f":"of:=[.G128]+273","v":286.1},34.9281,2.42,{"$f":"of:=([.J$135]*[.J128]*112/[.J$136])*([.H128]/273)","v":5.36948861019865},{"$f":"of:=[.K128]/[.M128]","v":{"v":0.719125305889092,"d":"72%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F128]+273.15))+143.3483*LN(([.F128]+273.15)/100)-21.8492*(([.F128]+273.15)/100)+[.I128]*(-0.033096+0.014259*(([.F128]+273.15)/100)-0.0017*((([.F128]+273.15)/100))^2))","v":7.46669400482308},null,5.3308]
[{"$nulls":4},3760,1.63,12.2,{"$f":"of:=[.G129]+273","v":285.2},34.8118,2.37,{"$f":"of:=([.J$135]*[.J129]*112/[.J$136])*([.H129]/273)","v":5.24200666610806},{"$f":"of:=[.K129]/[.M129]","v":{"v":0.679060491325808,"d":"68%"}},{"$f":"of:=EXP(-173.4292+249.6339*(100/([.F129]+273.15))+143.3483*LN(([.F129]+273.15)/100)-21.8492*(([.F129]+273.15)/100)+[.I129]*(-0.033096+0.014259*(([.F129]+273.15)/100)-0.0017*((([.F129]+273.15)/100))^2))","v":7.71949882678859},null,5.2519]
[]
[]
[]
[null,"DADOS COMPLEMENTARES"]
[]
[null,"FATOR DE CORREÇÃO TIOSSULFATO DE SÓDIO",{"$nulls":7},0.94517958]
[null,"VOLUME TOTAL DA ALICOTA UTILIZADA PARA ANÁLISE (ml)",{"$nulls":7},50]
[]
[null,"OBSERVAÇÕES GERAIS"]
[null," PERFIL Nº",null,null,3]
[null,"Nº ESTAÇÕES ",null,null,33]
[null,"REALIZADAS",null,null,15]
[null,"LEITURA DAS AMOSTRAS",null,null,90]
[]
[null,"FEITO POR : ",null,"Cb Hn Lourenço"]
[null,"CONFERIDO POR: ",null,"Cb Hn Ornelis"]

{"sheet":"PIR001","index":2,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":" PIR003","index":3,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR005","index":4,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR007","index":5,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR015","index":6,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR017","index":7,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR019","index":8,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR024","index":9,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR025","index":10,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR027","index":11,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR028","index":12,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR029","index":13,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR030","index":14,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR031","index":15,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"PIR033","index":16,"n_rows":0,"n_cols":0,"rows_kept":0}

{"sheet":"COMP. OD BANCADAxCTD","index":17,"n_rows":381,"n_cols":6,"rows_kept":381,"merged":[[0,0,0,5],[3,0,3,5],[339,3,339,5]]}
["PIRATA XIII / 2011."]
[]
[]
["COMPARAÇÃO ENTRE VALORES DE OD  CTD - BANCADA"]
[]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - BANCADA."]
["PIR-001",12,2,4.3296,4.2321,{"$f":"of:=[.D8]-[.E8]","v":0.0975000000000001}]
[null,10,79,4.3353,4.2146,{"$f":"of:=[.D9]-[.E9]","v":0.1207}]
[null,8,99,4.4961,4.2321,{"$f":"of:=[.D10]-[.E10]","v":0.264}]
[null,6,150,4.0513,4.0812,{"$f":"of:=[.D11]-[.E11]","v":-0.0298999999999996}]
[null,4,598,3.7717,3.8479,{"$f":"of:=[.D12]-[.E12]","v":-0.0762}]
[null,2,1496,5.0314,4.7608,{"$f":"of:=[.D13]-[.E13]","v":0.2706}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-003",12,3,4.3295,4.48882131531314,{"$f":"of:=[.D16]-[.E16]","v":-0.159321315313139}]
[null,10,108,4.0266,4.20752693817692,{"$f":"of:=[.D17]-[.E17]","v":-0.18092693817692}]
[null,8,128,3.7935,3.65927001458215,{"$f":"of:=[.D18]-[.E18]","v":0.13422998541785}]
[null,6,400,2.9854,2.99535261163126,{"$f":"of:=[.D19]-[.E19]","v":-0.0099526116312596}]
[null,4,685,3.3149,3.35613008085465,{"$f":"of:=[.D20]-[.E20]","v":-0.0412300808546497}]
[null,2,1493,5.1594,5.0098046410432,{"$f":"of:=[.D21]-[.E21]","v":0.1495953589568}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-005",12,3,4.3225,4.31281079986708,{"$f":"of:=[.D24]-[.E24]","v":0.00968920013291985}]
[null,10,95,3.9147,3.91389266869957,{"$f":"of:=[.D25]-[.E25]","v":0.000807331300429759}]
[null,8,113,2.058,2.04237118008271,{"$f":"of:=[.D26]-[.E26]","v":0.01562881991729}]
[null,6,495,2.8791,2.82304492007483,{"$f":"of:=[.D27]-[.E27]","v":0.0560550799251702}]
[null,4,698,3.322,3.35146137843692,{"$f":"of:=[.D28]-[.E28]","v":-0.02946137843692}]
[null,2,1470,5.0361,4.99442584221539,{"$f":"of:=[.D29]-[.E29]","v":0.0416741577846098}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-007",12,3,4.2945,4.27237456331225,{"$f":"of:=[.D32]-[.E32]","v":0.0221254366877499}]
[null,10,89,4.0834,4.06611408241231,{"$f":"of:=[.D33]-[.E33]","v":0.0172859175876905}]
[null,8,116,3.0899,3.04898840219766,{"$f":"of:=[.D34]-[.E34]","v":0.04091159780234}]
[null,6,399,3.1249,3.07162618019988,{"$f":"of:=[.D35]-[.E35]","v":0.0532738198001201}]
[null,4,700,2.9003,2.89350975424,{"$f":"of:=[.D36]-[.E36]","v":0.00679024575999998}]
[null,2,3512,5.705,5.760806528128,{"$f":"of:=[.D37]-[.E37]","v":-0.0558065281279996}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-015",12,4,4.2879,4.45515193637022,{"$f":"of:=[.D40]-[.E40]","v":-0.167251936370221}]
[null,10,98,2.888,2.85395762412308,{"$f":"of:=[.D41]-[.E41]","v":0.0340423758769202}]
[null,8,178,2.1909,2.29718771719089,{"$f":"of:=[.D42]-[.E42]","v":-0.10628771719089}]
[null,6,698,2.0995,2.18582132729305,{"$f":"of:=[.D43]-[.E43]","v":-0.0863213272930503}]
[null,4,1500,4.9432,4.9463893924224,{"$f":"of:=[.D44]-[.E44]","v":-0.00318939242240024}]
[null,2,4003,5.4623,5.25283309231262,{"$f":"of:=[.D45]-[.E45]","v":0.20946690768738}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-017",12,2,4.33,4.39424165599015,{"$f":"of:=[.D48]-[.E48]","v":-0.0642416559901502}]
[null,10,1,4.3351,4.37074303750892,{"$f":"of:=[.D49]-[.E49]","v":-0.0356430375089198}]
[null,8,17,4.3624,4.35343316409305,{"$f":"of:=[.D50]-[.E50]","v":0.00896683590694991}]
[null,6,142,2.1911,2.35006348643348,{"$f":"of:=[.D51]-[.E51]","v":-0.15896348643348}]
[null,4,597,1.8393,1.84167126335803,{"$f":"of:=[.D52]-[.E52]","v":-0.00237126335802973}]
[null,2,1498,4.7925,4.61377924409797,{"$f":"of:=[.D53]-[.E53]","v":0.17872075590203}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-019",12,1,4.352,4.58976877452111,{"$f":"of:=[.D56]-[.E56]","v":-0.237768774521109}]
[null,10,32,4.6497,4.54036738847311,{"$f":"of:=[.D57]-[.E57]","v":0.109332611526891}]
[null,8,55,4.7343,4.732955240864,{"$f":"of:=[.D58]-[.E58]","v":0.00134475913599985}]
[null,6,203,2.2023,2.53803692364406,{"$f":"of:=[.D59]-[.E59]","v":-0.33573692364406}]
[null,4,1502,4.6521,4.49606900274215,{"$f":"of:=[.D60]-[.E60]","v":0.15603099725785}]
[null,2,4080,5.4402,5.43467982701489,{"$f":"of:=[.D61]-[.E61]","v":0.00552017298510954}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-024",12,1,4.3444,4.498550705024,{"$f":"of:=[.D64]-[.E64]","v":-0.154150705024}]
[null,10,106,4.1326,4.2075866231168,{"$f":"of:=[.D65]-[.E65]","v":-0.0749866231167999}]
[null,8,501,2.9938,3.09190634103434,{"$f":"of:=[.D66]-[.E66]","v":-0.0981063410343399}]
[null,6,762,3.5222,3.676370494368,{"$f":"of:=[.D67]-[.E67]","v":-0.154170494368}]
[null,4,1507,5.1901,5.25283309231262,{"$f":"of:=[.D68]-[.E68]","v":-0.0627330923126195}]
[null,2,4010,5.5508,5.56444183507692,{"$f":"of:=[.D69]-[.E69]","v":-0.0136418350769203}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-025",12,1,4.3727,4.382608676544,{"$f":"of:=[.D72]-[.E72]","v":-0.00990867654399974}]
[null,10,93,4.2243,4.36941687785206,{"$f":"of:=[.D73]-[.E73]","v":-0.145116877852059}]
[null,8,706,3.2687,3.23638794575557,{"$f":"of:=[.D74]-[.E74]","v":0.03231205424443}]
[null,6,1111,3.837,3.90596673203692,{"$f":"of:=[.D75]-[.E75]","v":-0.0689667320369205}]
[null,4,2022,5.4172,5.39013327105575,{"$f":"of:=[.D76]-[.E76]","v":0.0270667289442503}]
[null,2,4060,5.343,5.28996556270474,{"$f":"of:=[.D77]-[.E77]","v":0.0530344372952598}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-027",12,1,4.4301,4.40432357151015,{"$f":"of:=[.D80]-[.E80]","v":0.0257764284898503}]
[null,10,81,4.4406,4.44770682952468,{"$f":"of:=[.D81]-[.E81]","v":-0.00710682952468034}]
[null,8,132,4.1153,4.15077890682142,{"$f":"of:=[.D82]-[.E82]","v":-0.0354789068214192}]
[null,6,308,2.205,2.30133681319335,{"$f":"of:=[.D83]-[.E83]","v":-0.09633681319335}]
[null,4,703,3.2631,3.26548590501022,{"$f":"of:=[.D84]-[.E84]","v":-0.00238590501021996}]
[null,2,1488,4.7985,4.78875476560738,{"$f":"of:=[.D85]-[.E85]","v":0.00974523439261965}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-028",12,6,4.4532,4.44621780815557,{"$f":"of:=[.D88]-[.E88]","v":0.00698219184442994}]
[null,10,134,4.3473,4.31568026813046,{"$f":"of:=[.D89]-[.E89]","v":0.0316197318695393}]
[null,8,306,2.5138,2.53440743405686,{"$f":"of:=[.D90]-[.E90]","v":-0.0206074340568603}]
[null,6,504,3.324,3.31480973786191,{"$f":"of:=[.D91]-[.E91]","v":0.00919026213808971}]
[null,4,709,3.4634,3.46197468317538,{"$f":"of:=[.D92]-[.E92]","v":0.0014253168246201}]
[null,2,1517,4.7825,4.86402789794363,{"$f":"of:=[.D93]-[.E93]","v":-0.08152789794363}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-029",12,6,4.4874,4.51417767408,{"$f":"of:=[.D96]-[.E96]","v":-0.0267776740799999}]
[null,10,120,4.4698,4.46338808581809,{"$f":"of:=[.D97]-[.E97]","v":0.00641191418191056}]
[null,8,710,3.8818,3.74294991339815,{"$f":"of:=[.D98]-[.E98]","v":0.13885008660185}]
[null,6,1107,4.0764,4.08026754073846,{"$f":"of:=[.D99]-[.E99]","v":-0.00386754073846074}]
[null,4,2019,5.5305,5.47350295708652,{"$f":"of:=[.D100]-[.E100]","v":0.0569970429134798}]
[null,2,4062,5.2941,5.28258249841625,{"$f":"of:=[.D101]-[.E101]","v":0.0115175015837501}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-030",12,1502,4.4745,4.57748434822597,{"$f":"of:=[.D104]-[.E104]","v":-0.10298434822597}]
[null,10,712,4.5146,4.505103950112,{"$f":"of:=[.D105]-[.E105]","v":0.0094960498879999}]
[null,8,533,3.9912,3.9972778653696,{"$f":"of:=[.D106]-[.E106]","v":-0.00607786536959987}]
[null,6,253,3.166,3.16458919661391,{"$f":"of:=[.D107]-[.E107]","v":0.00141080338609001}]
[null,3,115,4.0632,4.079143019392,{"$f":"of:=[.D108]-[.E108]","v":-0.015943019392}]
[null,2,1,4.9013,4.91035817741785,{"$f":"of:=[.D109]-[.E109]","v":-0.00905817741784976}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-031",12,1,4.5225,4.66839220493785,{"$f":"of:=[.D112]-[.E112]","v":-0.14589220493785}]
[null,10,161,4.4926,4.51756674875865,{"$f":"of:=[.D113]-[.E113]","v":-0.0249667487586498}]
[null,8,199,4.2721,4.2836042661376,{"$f":"of:=[.D114]-[.E114]","v":-0.0115042661376004}]
[null,6,449,3.4065,3.44094225633674,{"$f":"of:=[.D115]-[.E115]","v":-0.0344422563367401}]
[null,4,702,4.1837,4.19175026043077,{"$f":"of:=[.D116]-[.E116]","v":-0.00805026043077017}]
[null,2,1557,5.0429,5.01623380101711,{"$f":"of:=[.D117]-[.E117]","v":0.0266661989828902}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-033",12,1,4.5539,4.54978234650486,{"$f":"of:=[.D120]-[.E120]","v":0.00411765349514059}]
[null,10,138,4.4991,4.46846006485662,{"$f":"of:=[.D121]-[.E121]","v":0.0306399351433804}]
[null,8,507,4.0054,4.09463039269465,{"$f":"of:=[.D122]-[.E122]","v":-0.0892303926946498}]
[null,6,1008,4.1461,4.14282970419988,{"$f":"of:=[.D123]-[.E123]","v":0.00327029580012006}]
[null,4,2521,5.3308,5.36948861019865,{"$f":"of:=[.D124]-[.E124]","v":-0.0386886101986503}]
[null,2,3760,5.2519,5.24200666610806,{"$f":"of:=[.D125]-[.E125]","v":0.0098933338919398}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["PIR-033",12,{"$nulls":3},{"$f":"of:=[.D128]-[.E128]","v":0}]
[null,10,{"$nulls":3},{"$f":"of:=[.D129]-[.E129]","v":0}]
[null,8,{"$nulls":3},{"$f":"of:=[.D130]-[.E130]","v":0}]
[null,6,{"$nulls":3},{"$f":"of:=[.D131]-[.E131]","v":0}]
[null,4,{"$nulls":3},{"$f":"of:=[.D132]-[.E132]","v":0}]
[null,2,{"$nulls":3},{"$f":"of:=[.D133]-[.E133]","v":0}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 017",8,1,3.5602,3.91,{"$f":"of:=[.D136]-[.E136]","v":-0.3498}]
[null,7,40,3.5386,4.08,{"$f":"of:=[.D137]-[.E137]","v":-0.5414}]
[null,6,150,3.243,4.18,{"$f":"of:=[.D138]-[.E138]","v":-0.937}]
[null,5,400,3.0898,3.61,{"$f":"of:=[.D139]-[.E139]","v":-0.5202}]
[null,4,830,2.8102,4.28,{"$f":"of:=[.D140]-[.E140]","v":-1.4698}]
[null,3,1500,2.0852,3.36,{"$f":"of:=[.D141]-[.E141]","v":-1.2748}]
[null,2,2100,2.3101,4.01,{"$f":"of:=[.D142]-[.E142]","v":-1.6999}]
[null,1,2680,2.2787,3.56,{"$f":"of:=[.D143]-[.E143]","v":-1.2813}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 018",6,1,3.6056,3.64,{"$f":"of:=[.D146]-[.E146]","v":-0.0344000000000002}]
[null,5,50,3.5096,3.61,{"$f":"of:=[.D147]-[.E147]","v":-0.1004}]
[null,4,75,3.2195,3.73,{"$f":"of:=[.D148]-[.E148]","v":-0.5105}]
[null,3,95,2.9747,3.49,{"$f":"of:=[.D149]-[.E149]","v":-0.5153}]
[null,2,115,2.9203,3.54,{"$f":"of:=[.D150]-[.E150]","v":-0.6197}]
[null,1,125,2.9207,3.56,{"$f":"of:=[.D151]-[.E151]","v":-0.6393}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 019",6,2,3.5639,3.61,{"$f":"of:=[.D154]-[.E154]","v":-0.0460999999999996}]
[null,5,25,3.5654,3.64,{"$f":"of:=[.D155]-[.E155]","v":-0.0746000000000002}]
[null,4,50,3.5148,3.68,{"$f":"of:=[.D156]-[.E156]","v":-0.1652}]
[null,3,80,3.0855,3.46,{"$f":"of:=[.D157]-[.E157]","v":-0.3745}]
[null,2,110,2.969,3.39,{"$f":"of:=[.D158]-[.E158]","v":-0.421}]
[null,1,120,2.9661,3.41,{"$f":"of:=[.D159]-[.E159]","v":-0.4439}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 020",6,2,3.5491,3.71,{"$f":"of:=[.D162]-[.E162]","v":-0.1609}]
[null,5,20,3.5484,3.68,{"$f":"of:=[.D163]-[.E163]","v":-0.1316}]
[null,4,40,3.476,3.59,{"$f":"of:=[.D164]-[.E164]","v":-0.114}]
[null,3,60,3.3198,3.46,{"$f":"of:=[.D165]-[.E165]","v":-0.1402}]
[null,2,120,3.1358,3.49,{"$f":"of:=[.D166]-[.E166]","v":-0.3542}]
[null,1,148,3.007,3.34,{"$f":"of:=[.D167]-[.E167]","v":-0.333}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 021",8,1,3.6393,3.73,{"$f":"of:=[.D170]-[.E170]","v":-0.0907}]
[null,7,20,3.6187,3.71,{"$f":"of:=[.D171]-[.E171]","v":-0.0912999999999999}]
[null,6,120,3.4835,3.51,{"$f":"of:=[.D172]-[.E172]","v":-0.0264999999999995}]
[null,5,400,2.9527,3.36,{"$f":"of:=[.D173]-[.E173]","v":-0.4073}]
[null,4,680,2.5839,3.21,{"$f":"of:=[.D174]-[.E174]","v":-0.6261}]
[null,3,780,2.4926,3.12,{"$f":"of:=[.D175]-[.E175]","v":-0.6274}]
[null,2,1464,2.1886,3.14,{"$f":"of:=[.D176]-[.E176]","v":-0.9514}]
[null,1,2148,2.3659,3.07,{"$f":"of:=[.D177]-[.E177]","v":-0.7041}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 022",8,1,3.6039,2.68,{"$f":"of:=[.D180]-[.E180]","v":0.9239}]
[null,7,15,3.6058,3.71,{"$f":"of:=[.D181]-[.E181]","v":-0.1042}]
[null,6,85,3.3205,3.41,{"$f":"of:=[.D182]-[.E182]","v":-0.0895000000000001}]
[null,5,350,3.0574,3.51,{"$f":"of:=[.D183]-[.E183]","v":-0.4526}]
[null,4,740,2.7332,3.17,{"$f":"of:=[.D184]-[.E184]","v":-0.4368}]
[null,3,900,2.5746,3.24,{"$f":"of:=[.D185]-[.E185]","v":-0.6654}]
[null,2,1740,2.1881,3.19,{"$f":"of:=[.D186]-[.E186]","v":-1.0019}]
[null,1,3100,2.172,2.99,{"$f":"of:=[.D187]-[.E187]","v":-0.818}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 023",8,1,3.578,3.78,{"$f":"of:=[.D190]-[.E190]","v":-0.202}]
[null,7,20,3.5947,3.81,{"$f":"of:=[.D191]-[.E191]","v":-0.2153}]
[null,6,40,3.6085,3.91,{"$f":"of:=[.D192]-[.E192]","v":-0.3015}]
[null,5,250,3.14,3.51,{"$f":"of:=[.D193]-[.E193]","v":-0.37}]
[null,4,750,2.5586,3.21,{"$f":"of:=[.D194]-[.E194]","v":-0.6514}]
[null,3,970,2.3638,3.14,{"$f":"of:=[.D195]-[.E195]","v":-0.7762}]
[null,2,1800,2.2144,3.04,{"$f":"of:=[.D196]-[.E196]","v":-0.8256}]
[null,1,2730,2.2018,3.07,{"$f":"of:=[.D197]-[.E197]","v":-0.8682}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 024",6,1,3.5428,3.59,{"$f":"of:=[.D200]-[.E200]","v":-0.0471999999999997}]
[null,5,25,3.5581,3.71,{"$f":"of:=[.D201]-[.E201]","v":-0.1519}]
[null,4,40,3.5607,3.56,{"$f":"of:=[.D202]-[.E202]","v":0.000699999999999701}]
[null,3,290,3.1414,3.51,{"$f":"of:=[.D203]-[.E203]","v":-0.3686}]
[null,2,400,2.9818,3.39,{"$f":"of:=[.D204]-[.E204]","v":-0.4082}]
[null,1,600,2.6052,3.26,{"$f":"of:=[.D205]-[.E205]","v":-0.6548}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 025",6,1,3.5594,3.83,{"$f":"of:=[.D208]-[.E208]","v":-0.2706}]
[null,5,20,3.555,3.81,{"$f":"of:=[.D209]-[.E209]","v":-0.255}]
[null,4,50,3.5565,3.81,{"$f":"of:=[.D210]-[.E210]","v":-0.2535}]
[null,3,130,3.0139,3.54,{"$f":"of:=[.D211]-[.E211]","v":-0.5261}]
[null,2,170,2.9465,3.12,{"$f":"of:=[.D212]-[.E212]","v":-0.1735}]
[null,1,195,2.9142,3.09,{"$f":"of:=[.D213]-[.E213]","v":-0.1758}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 026",6,1,3.5643,3.59,{"$f":"of:=[.D216]-[.E216]","v":-0.0256999999999996}]
[null,5,20,3.5662,3.61,{"$f":"of:=[.D217]-[.E217]","v":-0.0437999999999996}]
[null,4,40,3.5523,3.46,{"$f":"of:=[.D218]-[.E218]","v":0.0922999999999998}]
[null,3,70,3.2043,3.34,{"$f":"of:=[.D219]-[.E219]","v":-0.1357}]
[null,2,90,3.0115,3.19,{"$f":"of:=[.D220]-[.E220]","v":-0.1785}]
[null,1,120,2.924,3.04,{"$f":"of:=[.D221]-[.E221]","v":-0.116}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 027",6,1,3.5574,3.73,{"$f":"of:=[.D224]-[.E224]","v":-0.1726}]
[null,5,20,3.598,3.64,{"$f":"of:=[.D225]-[.E225]","v":-0.0420000000000003}]
[null,4,30,3.5934,3.61,{"$f":"of:=[.D226]-[.E226]","v":-0.0165999999999999}]
[null,3,45,3.3196,3.61,{"$f":"of:=[.D227]-[.E227]","v":-0.2904}]
[null,2,75,3.0096,3.24,{"$f":"of:=[.D228]-[.E228]","v":-0.2304}]
[null,1,110,2.9029,2.94,{"$f":"of:=[.D229]-[.E229]","v":-0.0371000000000001}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 028",4,1,3.6159,3.68,{"$f":"of:=[.D232]-[.E232]","v":-0.0641000000000003}]
[null,3,10,3.6225,3.68,{"$f":"of:=[.D233]-[.E233]","v":-0.0575000000000001}]
[null,2,35,3.3587,3.39,{"$f":"of:=[.D234]-[.E234]","v":-0.0313000000000003}]
[null,1,80,2.8155,3.17,{"$f":"of:=[.D235]-[.E235]","v":-0.3545}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 029",4,2.5,3.6277,3.76,{"$f":"of:=[.D238]-[.E238]","v":-0.1323}]
[null,3,15,3.5132,3.71,{"$f":"of:=[.D239]-[.E239]","v":-0.1968}]
[null,2,30,2.9277,3.51,{"$f":"of:=[.D240]-[.E240]","v":-0.5823}]
[null,1,47,2.5357,3.39,{"$f":"of:=[.D241]-[.E241]","v":-0.8543}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 030",4,2,3.6684,3.81,{"$f":"of:=[.D244]-[.E244]","v":-0.1416}]
[null,3,7,3.6664,3.83,{"$f":"of:=[.D245]-[.E245]","v":-0.1636}]
[null,2,14,2.6546,3.51,{"$f":"of:=[.D246]-[.E246]","v":-0.8554}]
[null,1,20,2.5682,3.39,{"$f":"of:=[.D247]-[.E247]","v":-0.8218}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 031",4,1,3.6779,3.66,{"$f":"of:=[.D250]-[.E250]","v":0.0179}]
[null,3,10,3.0962,3.44,{"$f":"of:=[.D251]-[.E251]","v":-0.3438}]
[null,2,25,2.5958,3.39,{"$f":"of:=[.D252]-[.E252]","v":-0.7942}]
[null,1,40,2.5919,3.31,{"$f":"of:=[.D253]-[.E253]","v":-0.7181}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 032",4,1,3.577,3.61,{"$f":"of:=[.D256]-[.E256]","v":-0.0329999999999999}]
[null,3,10,3.5917,3.41,{"$f":"of:=[.D257]-[.E257]","v":0.1817}]
[null,2,40,3.1095,3.26,{"$f":"of:=[.D258]-[.E258]","v":-0.1505}]
[null,1,75,2.8524,3.14,{"$f":"of:=[.D259]-[.E259]","v":-0.2876}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 033",6,1,3.5588,3.64,{"$f":"of:=[.D262]-[.E262]","v":-0.0811999999999999}]
[null,5,10,3.5753,3.61,{"$f":"of:=[.D263]-[.E263]","v":-0.0347}]
[null,4,20,3.5829,3.49,{"$f":"of:=[.D264]-[.E264]","v":0.0928999999999998}]
[null,3,40,3.5202,3.41,{"$f":"of:=[.D265]-[.E265]","v":0.1102}]
[null,2,80,3.0208,3.12,{"$f":"of:=[.D266]-[.E266]","v":-0.0992000000000002}]
[null,1,110,2.9522,2.99,{"$f":"of:=[.D267]-[.E267]","v":-0.0378000000000003}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 034",6,1,3.5437,3.54,{"$f":"of:=[.D270]-[.E270]","v":0.00370000000000026}]
[null,5,10,3.56,3.5505,{"$f":"of:=[.D271]-[.E271]","v":0.00950000000000006}]
[null,4,30,3.51,3.5911,{"$f":"of:=[.D272]-[.E272]","v":-0.0811000000000002}]
[null,3,50,3.39,3.3893,{"$f":"of:=[.D273]-[.E273]","v":0.000700000000000145}]
[null,2,125,2.97,2.9232,{"$f":"of:=[.D274]-[.E274]","v":0.0467999999999997}]
[null,1,200,2.94,2.8116,{"$f":"of:=[.D275]-[.E275]","v":0.1284}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 035",8,2,3.5319,3.59,{"$f":"of:=[.D278]-[.E278]","v":-0.0580999999999996}]
[null,7,100,3.2044,3.51,{"$f":"of:=[.D279]-[.E279]","v":-0.3056}]
[null,6,150,3.1305,3.39,{"$f":"of:=[.D280]-[.E280]","v":-0.2595}]
[null,5,350,2.7401,3.21,{"$f":"of:=[.D281]-[.E281]","v":-0.4699}]
[null,4,752,2.508,3.09,{"$f":"of:=[.D282]-[.E282]","v":-0.582}]
[null,3,899,2.3922,3.02,{"$f":"of:=[.D283]-[.E283]","v":-0.6278}]
[null,2,1090,2.2248,2.99,{"$f":"of:=[.D284]-[.E284]","v":-0.7652}]
[null,1,1200,2.1756,2.99,{"$f":"of:=[.D285]-[.E285]","v":-0.8144}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 036",8,1,3.5559,3.71,{"$f":"of:=[.D288]-[.E288]","v":-0.1541}]
[null,7,20,3.5762,3.73,{"$f":"of:=[.D289]-[.E289]","v":-0.1538}]
[null,6,50,3.5602,3.66,{"$f":"of:=[.D290]-[.E290]","v":-0.0998000000000001}]
[null,5,350,2.9977,3.39,{"$f":"of:=[.D291]-[.E291]","v":-0.3923}]
[null,4,750,2.569,3.24,{"$f":"of:=[.D292]-[.E292]","v":-0.671}]
[null,3,1000,2.3565,3.07,{"$f":"of:=[.D293]-[.E293]","v":-0.7135}]
[null,2,1450,2.1388,2.97,{"$f":"of:=[.D294]-[.E294]","v":-0.8312}]
[null,1,1875,2.3046,2.99,{"$f":"of:=[.D295]-[.E295]","v":-0.6854}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 037",8,1,3.5815,3.81,{"$f":"of:=[.D298]-[.E298]","v":-0.2285}]
[null,7,10,3.5758,3.66,{"$f":"of:=[.D299]-[.E299]","v":-0.0842000000000001}]
[null,6,37,3.6461,3.88,{"$f":"of:=[.D300]-[.E300]","v":-0.2339}]
[null,5,300,3.0179,3.51,{"$f":"of:=[.D301]-[.E301]","v":-0.4921}]
[null,4,780,2.6494,3.04,{"$f":"of:=[.D302]-[.E302]","v":-0.3906}]
[null,3,1020,2.3992,3.12,{"$f":"of:=[.D303]-[.E303]","v":-0.7208}]
[null,2,1570,2.1698,2.92,{"$f":"of:=[.D304]-[.E304]","v":-0.7502}]
[null,1,2200,2.3676,2.89,{"$f":"of:=[.D305]-[.E305]","v":-0.5224}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 038",8,1,3.5721,3.71,{"$f":"of:=[.D308]-[.E308]","v":-0.1379}]
[null,7,25,3.5892,3.71,{"$f":"of:=[.D309]-[.E309]","v":-0.1208}]
[null,6,150,3.1743,3.26,{"$f":"of:=[.D310]-[.E310]","v":-0.0856999999999997}]
[null,5,350,2.9576,3.09,{"$f":"of:=[.D311]-[.E311]","v":-0.1324}]
[null,4,800,2.5761,3.02,{"$f":"of:=[.D312]-[.E312]","v":-0.4439}]
[null,3,1000,2.4001,2.99,{"$f":"of:=[.D313]-[.E313]","v":-0.5899}]
[null,2,1500,2.1141,2.89,{"$f":"of:=[.D314]-[.E314]","v":-0.7759}]
[null,1,2800,2.2234,2.94,{"$f":"of:=[.D315]-[.E315]","v":-0.7166}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 039",8,2,3.59,3.76,{"$f":"of:=[.D318]-[.E318]","v":-0.17}]
[null,7,30,3.6005,3.76,{"$f":"of:=[.D319]-[.E319]","v":-0.1595}]
[null,6,150,3.042,3.54,{"$f":"of:=[.D320]-[.E320]","v":-0.498}]
[null,5,650,2.5659,3.09,{"$f":"of:=[.D321]-[.E321]","v":-0.5241}]
[null,4,800,2.5793,3.12,{"$f":"of:=[.D322]-[.E322]","v":-0.5407}]
[null,3,1350,2.1211,2.89,{"$f":"of:=[.D323]-[.E323]","v":-0.7689}]
[null,2,2000,2.3498,2.89,{"$f":"of:=[.D324]-[.E324]","v":-0.5402}]
[null,1,3000,2.1714,2.82,{"$f":"of:=[.D325]-[.E325]","v":-0.6486}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 040",8,2,3.6065,3.76,{"$f":"of:=[.D328]-[.E328]","v":-0.1535}]
[null,7,42,3.6631,3.78,{"$f":"of:=[.D329]-[.E329]","v":-0.1169}]
[null,6,49,3.6188,3.78,{"$f":"of:=[.D330]-[.E330]","v":-0.1612}]
[null,5,244,3.0247,3.09,{"$f":"of:=[.D331]-[.E331]","v":-0.0652999999999997}]
[null,4,745,2.6103,2.97,{"$f":"of:=[.D332]-[.E332]","v":-0.3597}]
[null,3,844,2.5243,2.89,{"$f":"of:=[.D333]-[.E333]","v":-0.3657}]
[null,2,1498,2.1327,2.77,{"$f":"of:=[.D334]-[.E334]","v":-0.6373}]
[null,1,2403,2.342,2.84,{"$f":"of:=[.D335]-[.E335]","v":-0.498}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 041",8,1,3.5686,3.78,{"$f":"of:=[.D338]-[.E338]","v":-0.2114}]
[null,7,10,3.5802,3.76,{"$f":"of:=[.D339]-[.E339]","v":-0.1798}]
[null,6,40,"NÃO FECHOU"]
[null,5,200,3.0827,3.46,{"$f":"of:=[.D341]-[.E341]","v":-0.3773}]
[null,4,800,2.5591,3.24,{"$f":"of:=[.D342]-[.E342]","v":-0.6809}]
[null,3,1200,2.1731,3.09,{"$f":"of:=[.D343]-[.E343]","v":-0.9169}]
[null,2,1500,2.1611,3.07,{"$f":"of:=[.D344]-[.E344]","v":-0.9089}]
[null,1,2000,2.3554,2.92,{"$f":"of:=[.D345]-[.E345]","v":-0.5646}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 042",6,1,3.5731,3.71,{"$f":"of:=[.D348]-[.E348]","v":-0.1369}]
[null,5,20,3.5768,3.73,{"$f":"of:=[.D349]-[.E349]","v":-0.1532}]
[null,4,40,3.6119,3.59,{"$f":"of:=[.D350]-[.E350]","v":0.0219}]
[null,3,150,3.2208,3.31,{"$f":"of:=[.D351]-[.E351]","v":-0.0891999999999999}]
[null,2,300,2.9022,3.02,{"$f":"of:=[.D352]-[.E352]","v":-0.1178}]
[null,1,450,2.7758,2.92,{"$f":"of:=[.D353]-[.E353]","v":-0.1442}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 043",6,2,3.5359,3.66,{"$f":"of:=[.D356]-[.E356]","v":-0.1241}]
[null,5,20,3.5951,3.68,{"$f":"of:=[.D357]-[.E357]","v":-0.0849000000000002}]
[null,4,45,3.38,3.44,{"$f":"of:=[.D358]-[.E358]","v":-0.0600000000000001}]
[null,3,100,2.9879,3.12,{"$f":"of:=[.D359]-[.E359]","v":-0.1321}]
[null,2,148,3.0001,3.17,{"$f":"of:=[.D360]-[.E360]","v":-0.1699}]
[null,1,190,2.9155,3.02,{"$f":"of:=[.D361]-[.E361]","v":-0.1045}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 044",6,2,3.5821,3.76,{"$f":"of:=[.D364]-[.E364]","v":-0.1779}]
[null,5,9,3.5772,3.73,{"$f":"of:=[.D365]-[.E365]","v":-0.1528}]
[null,4,20,3.6775,3.86,{"$f":"of:=[.D366]-[.E366]","v":-0.1825}]
[null,3,30,3.6395,3.83,{"$f":"of:=[.D367]-[.E367]","v":-0.1905}]
[null,2,100,2.9245,3.04,{"$f":"of:=[.D368]-[.E368]","v":-0.1155}]
[null,1,120,2.8786,2.99,{"$f":"of:=[.D369]-[.E369]","v":-0.1114}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 045",4,1,3.5739,3.39,{"$f":"of:=[.D372]-[.E372]","v":0.1839}]
[null,3,14,3.5681,3.54,{"$f":"of:=[.D373]-[.E373]","v":0.0281000000000002}]
[null,2,30,3.5031,3.49,{"$f":"of:=[.D374]-[.E374]","v":0.0130999999999997}]
[null,1,100,2.8795,3.73,{"$f":"of:=[.D375]-[.E375]","v":-0.8505}]
["ESTAÇÃO","GARRAFA","PROFUNDIDADE","OD","OD","DIFERENÇA"]
[null,"Nº","COLETA (m)","CTD","BANCADA","CTD - POR"]
["OCS-4 046",4,2,3.5816,3.76,{"$f":"of:=[.D378]-[.E378]","v":-0.1784}]
[null,3,20,3.563,3.73,{"$f":"of:=[.D379]-[.E379]","v":-0.167}]
[null,2,40,2.8393,3.04,{"$f":"of:=[.D380]-[.E380]","v":-0.2007}]
[null,1,50,2.716,2.94,{"$f":"of:=[.D381]-[.E381]","v":-0.224}]